## General
S1 INTRAWELL ONLY
ROC CURVE SUMMARY
Sample size | 446
Positive group a | 62 (13.90%)
Negative group b | 384 (86.10%)
Sample sets included in ROC: PAMF Pre-COVID (311), Loma Linda PCR+ (27), Loma Linda PCR- (29), FL PCR+ (35), CDPH PCR- (4), In-house healthy (40)
All sample data: Above and cross-reactivity (73), CDPH PCR+ (41), PAMF PCR+ (33)
Final cutoffs (MedCalc ROC curve analysis, adj*) | [5 empty] | PCR+ with known days since symptom onset data, overview
| Criterion (≤) | Sensitivity | Specificity |  |  | Days since symptom onset | # samples | IgG+ | IgM+ | Ab+ | % IgG+ | % IgM+ | % Ab+
IgM | 1.38 | 70.9677 | 99.74 |  |  | 0 to 7 | 16 | 2 | 7 | 7 | 12.5% | 43.75% | 43.75%
IgG | 1.62 | 51.6129 | 99.74 |  |  | 8 to 14 | 21 | 10 | 14 | 14 | 47.62% | 66.67% | 66.67%
1 IgG false pos: PAMF Pre-COVID 434 | [5 empty] | > 14 | 33 | 33 | 33 | 33 | 100% | 100% | 100%
1 IgM false pos: PAMF Pre-COVID 434
*IgM: intrawell ratio adjusted by a factor of 100. IgG: intrawell ratio adjusted by a factor of 10.
Intrawell Standard MFI
Avg | 17714.4739
SD | 4156.2194
| Criterion | Sensitivity | 95% CI | Specificity | 95% CI |  | Criterion | Sensitivity | 95% CI | Specificity | 95% CI
| 1.7382 | 50 | 37.0 - 63.0 | 99.74 | 98.6 - 100.0 |  | 1.3614 | 70.97 | 58.1 - 81.8 | 99.74 | 98.6 - 100.0

## ROC Curve analysis
Sheet_Order | Source/Sample Type | Expected Result | Patient ID | PCR+: Days Since Symptom Onset | IgM MFI | IgG MFI | Intrawell Standard MFI | IgM Intrawell Ratio | IgG Intrawell Ratio | IgM Intrawell, adjusted | IgG Intrawell, adjusted | Status | S1 Intrawell_IgM (<=1.38) | S1 Intrawell_IgG (<=1.62)
34 | CDPH PCR- | Negative | 2.1 |  | 2.3333 | 99 | 15217.6667 | -0.0034171 | 0.0015771 | -0.34171 | 0.015771 | H | neg | neg
35 | CDPH PCR- | Negative | 2.2 |  | -9 | 194.3333 | 18417.5 | -0.0040994 | 0.0063889 | -0.40994 | 0.063889 | H | neg | neg
36 | CDPH PCR- | Negative | 2.3 |  | -4.6667 | 186.3333 | 17667.3333 | -0.0040281 | 0.0062073 | -0.40281 | 0.062073 | H | neg | neg
37 | CDPH PCR- | Negative | 2.4 |  | -1.3333 | 106.6667 | 16914.6667 | -0.0040103 | 0.0017736 | -0.40103 | 0.017736 | H | neg | neg
79 | LL PCR- | Negative | 06872ER |  | 0.16667 | 146.1667 | 17672.6667 | -0.0037534 | 0.0039326 | -0.37534 | 0.039326 | H | neg | neg
80 | LL PCR- | Negative | 98671DR |  | 3.6667 | 1135.3333 | 16660.6667 | -0.0037714 | 0.063543 | -0.37714 | 0.63543 | H | neg | neg
81 | LL PCR- | Negative | 68409DP |  | -0.5 | 95.8333 | 16952.6667 | -0.0039522 | 0.0011306 | -0.39522 | 0.011306 | H | neg | neg
82 | LL PCR- | Negative | 52104RH |  | 3 | 95.6667 | 17116 | -0.00371 | 0.0011101 | -0.371 | 0.011101 | H | neg | neg
83 | LL PCR- | Negative | 33492KM |  | -3.3333 | 94 | 16197 | -0.0043115 | 0.0010702 | -0.43115 | 0.010702 | H | neg | neg
84 | LL PCR- | Negative | 73605JA |  | -25.8333 | 207.8333 | 16552 | -0.0055784 | 0.0079245 | -0.55784 | 0.079245 | H | neg | neg
85 | LL PCR- | Negative | 27215JD |  | 5 | 95 | 16476.6667 | -0.0037326 | 0.0011127 | -0.37326 | 0.011127 | H | neg | neg
86 | LL PCR- | Negative | 27829JP |  | 41.6667 | 113.8333 | 16551 | -0.0015004 | 0.0022456 | -0.15004 | 0.022456 | H | neg | neg
87 | LL PCR- | Negative | 08751EC |  | -1 | 97.3333 | 17456.3333 | -0.0038668 | 0.0011839 | -0.38668 | 0.011839 | H | neg | neg
88 | LL PCR- | Negative | 05458WP |  | 2.3333 | 99.3333 | 16660.5 | -0.0038514 | 0.0013605 | -0.38514 | 0.013605 | H | neg | neg
89 | LL PCR- | Negative | 45500BH |  | 1.6667 | 449.1667 | 17867 | -0.0036287 | 0.020848 | -0.36287 | 0.20848 | H | neg | neg
90 | LL PCR- | Negative | 09742JA |  | 6 | 168.6667 | 18620.8333 | 0.0014321 | 0.0049228 | 0.14321 | 0.049228 | H | neg | neg
91 | LL PCR- | Negative | 20443PM |  | -718.6667 | 133 | 18367.1667 | -0.038003 | 0.0030489 | -3.8003 | 0.030489 | H | neg | neg
92 | LL PCR- | Negative | 66311AC |  | 2.6667 | 67.8333 | 16714.5 | 0.00019943 | -0.00036894 | 0.019943 | -0.0036894 | H | neg | neg
93 | LL PCR- | Negative | 42089SR |  | 99.1667 | 136.8333 | 15572.3333 | 0.0064109 | 0.0040349 | 0.64109 | 0.040349 | H | neg | neg
94 | LL PCR- | Negative | 02065SJ |  | 24 | 67.1667 | 17770 | 0.0011255 | -0.00045958 | 0.11255 | -0.0045958 | H | neg | neg
95 | LL PCR- | Negative | 46577ME |  | 6 | 103 | 17217.6667 | 0.00011616 | 0.0016069 | 0.011616 | 0.016069 | H | neg | neg
96 | LL PCR- | Negative | 07018AA |  | 63.5 | 109.6667 | 17985 | 0.0033824 | 0.002984 | 0.33824 | 0.02984 | H | neg | neg
97 | LL PCR- | Negative | 56891ER |  | -0.66667 | 72.3333 | 17840 | -0.00018685 | 0.00091555 | -0.018685 | 0.0091555 | H | neg | neg
98 | LL PCR- | Negative | 42334BS |  | -2.5 | 191.3333 | 19103 | -0.00027046 | 0.0070844 | -0.027046 | 0.070844 | H | neg | neg
99 | LL PCR- | Negative | 60812KW |  | 76.3333 | 106.5 | 18412 | 0.004001 | 0.0027428 | 0.4001 | 0.027428 | H | neg | neg
100 | LL PCR- | Negative | 54431LR |  | 77.5 | 66.6667 | 15936.8333 | 0.00049152 | -0.00064839 | 0.049152 | -0.0064839 | H | neg | neg
101 | LL PCR- | Negative | 30056LC |  | 204.3333 | 91.3333 | 15164.3333 | 0.0088805 | 0.0009452 | 0.88805 | 0.009452 | H | neg | neg
102 | LL PCR- | Negative | 76896SA |  | 3.3333 | 69 | 19046.5 | 3.5002e-05 | 0.00068254 | 0.0035002 | 0.0068254 | H | neg | neg
103 | LL PCR- | Negative | 28570MR |  | 95.6667 | 108.1667 | 17212.5 | 0.0054031 | 0.0030307 | 0.54031 | 0.030307 | H | neg | neg
104 | LL PCR- | Negative | 24087RS |  | 35.6667 | 34.3333 | 17824.3333 | 0.0018514 | -0.0012156 | 0.18514 | -0.012156 | H | neg | neg
105 | LL PCR- | Negative | 25431LY |  | 56 | 7.6667 | 16335.5 | 0.0032649 | -0.0029588 | 0.32649 | -0.029588 | H | neg | neg
106 | LL PCR- | Negative | 08286QA |  | -5.3333 | 103 | 16198.1667 | -0.00049388 | 0.0029016 | -0.049388 | 0.029016 | H | neg | neg
107 | LL PCR- | Negative | 62379AD |  | 105.6667 | 105.8333 | 16364.6667 | 0.006294 | 0.0030452 | 0.6294 | 0.030452 | H | neg | neg
108 | PAMF Pre-COVID | Negative | 2 |  | 127.8333 | 102.3333 | 13574.3333 | 0.0068634 | 0.0035361 | 0.68634 | 0.035361 | H | neg | neg
109 | PAMF Pre-COVID | Negative | 8 |  | 2.3333 | 204.6667 | 12637.1667 | -0.0025586 | 0.011896 | -0.25586 | 0.11896 | H | neg | neg
110 | PAMF Pre-COVID | Negative | 13 |  | 103 | 191 | 13037.1667 | 0.0052414 | 0.010483 | 0.52414 | 0.10483 | H | neg | neg
111 | PAMF Pre-COVID | Negative | 14 |  | 24 | 113.1667 | 15773.1667 | -0.00038039 | 0.001511 | -0.038039 | 0.01511 | H | neg | neg
112 | PAMF Pre-COVID | Negative | 18 |  | -0.66667 | 78.3333 | 15728.5 | -0.0019498 | -0.00069937 | -0.19498 | -0.0069937 | H | neg | neg
113 | PAMF Pre-COVID | Negative | 22 |  | -1.8333 | 85 | 17106 | -0.0018609 | -0.00025332 | -0.18609 | -0.0025332 | H | neg | neg
114 | PAMF Pre-COVID | Negative | 23 |  | 2.3333 | 117.3333 | 15504 | -0.0017845 | 0.001806 | -0.17845 | 0.01806 | H | neg | neg
115 | PAMF Pre-COVID | Negative | 29 |  | 34.8333 | 143.8333 | 15466.1667 | 0.00031251 | 0.0035238 | 0.031251 | 0.035238 | H | neg | neg
116 | PAMF Pre-COVID | Negative | 33 |  | 103.8333 | 188 | 14617.1667 | 0.0050511 | 0.0067501 | 0.50511 | 0.067501 | H | neg | neg
117 | PAMF Pre-COVID | Negative | 81 |  | 31.3333 | 105.8333 | 16307.6667 | 8.1761e-05 | 0.0010118 | 0.0081761 | 0.010118 | H | neg | neg
118 | PAMF Pre-COVID | Negative | 84 |  | 37 | 147.5 | 15768 | 0.00044394 | 0.0036889 | 0.044394 | 0.036889 | H | neg | neg
119 | PAMF Pre-COVID | Negative | 87 |  | 5.3333 | 111.3333 | 15810.1667 | -0.0015602 | 0.0013915 | -0.15602 | 0.013915 | H | neg | neg
120 | PAMF Pre-COVID | Negative | 88 |  | 2 | 104.3333 | 15570.6667 | -0.0017983 | 0.00096335 | -0.17983 | 0.0096335 | H | neg | neg
121 | PAMF Pre-COVID | Negative | 200 |  | -104.67 | 35.67 | 21178.33 | -0.00512 | 0.00022 | -0.512 | 0.0022 | H | neg | neg
122 | PAMF Pre-COVID | Negative | 201 |  | -0.83 | 41.17 | 16851.5 | -0.00027 | 0.0006 | -0.027 | 0.006 | H | neg | neg
123 | PAMF Pre-COVID | Negative | 202 |  | -42.17 | 65.33 | 17560.17 | -0.00261 | 0.00196 | -0.261 | 0.0196 | H | neg | neg
124 | PAMF Pre-COVID | Negative | 203 |  | -34 | 101.5 | 16089.83 | -0.00234 | 0.00438 | -0.234 | 0.0438 | H | neg | neg
125 | PAMF Pre-COVID | Negative | 204 |  | -1.17 | 543 | 17505.5 | -0.00028 | 0.02925 | -0.028 | 0.2925 | H | neg | neg
126 | PAMF Pre-COVID | Negative | 205 |  | 5 | 98.33 | 15260.17 | 9e-05 | 0.00441 | 0.009 | 0.0441 | H | neg | neg
127 | PAMF Pre-COVID | Negative | 206 |  | 1.33 | 96.67 | 16554 | -0.00014 | 0.00397 | -0.014 | 0.0397 | H | neg | neg
128 | PAMF Pre-COVID | Negative | 207 |  | -5 | 94 | 16982.67 | -0.00051 | 0.00371 | -0.051 | 0.0371 | H | neg | neg
129 | PAMF Pre-COVID | Negative | 208 |  | 0.83 | 89.83 | 16320.67 | -0.00017 | 0.0036 | -0.017 | 0.036 | H | neg | neg
130 | PAMF Pre-COVID | Negative | 209 |  | -1.5 | 91.33 | 17744.5 | -0.00029 | 0.0034 | -0.029 | 0.034 | H | neg | neg
131 | PAMF Pre-COVID | Negative | 210 |  | -2.33 | 97.33 | 16194.67 | -0.00037 | 0.0041 | -0.037 | 0.041 | H | neg | neg
132 | PAMF Pre-COVID | Negative | 211 |  | -1.33 | 97.33 | 15641.83 | -0.00032 | 0.00424 | -0.032 | 0.0424 | H | neg | neg
133 | PAMF Pre-COVID | Negative | 212 |  | -106.33 | 6 | 22540.33 | -0.006 | -0.00314 | -0.6 | -0.0314 | H | neg | neg
134 | PAMF Pre-COVID | Negative | 213 |  | 5 | 112.17 | 21591.5 | -0.0011 | 0.00164 | -0.11 | 0.0164 | H | neg | neg
135 | PAMF Pre-COVID | Negative | 214 |  | -392 | 29.17 | 20342 | -0.02069 | -0.00234 | -2.069 | -0.0234 | H | neg | neg
136 | PAMF Pre-COVID | Negative | 215 |  | -27.33 | 43 | 20119.5 | -0.00279 | -0.00167 | -0.279 | -0.0167 | H | neg | neg
137 | PAMF Pre-COVID | Negative | 216 |  | 3.33 | 101.67 | 19878.33 | -0.00128 | 0.00126 | -0.128 | 0.0126 | H | neg | neg
138 | PAMF Pre-COVID | Negative | 217 |  | 70.33 | 101.33 | 19927.5 | 0.00208 | 0.00124 | 0.208 | 0.0124 | H | neg | neg
139 | PAMF Pre-COVID | Negative | 218 |  | -4.33 | 9.5 | 20779.5 | -0.0016 | -0.00323 | -0.16 | -0.0323 | H | neg | neg
140 | PAMF Pre-COVID | Negative | 219 |  | -3.33 | 97.67 | 21724.17 | -0.00148 | 0.00097 | -0.148 | 0.0097 | H | neg | neg
141 | PAMF Pre-COVID | Negative | 220 |  | -2.33 | 16.5 | 20032.67 | -0.00156 | -0.003 | -0.156 | -0.03 | H | neg | neg
142 | PAMF Pre-COVID | Negative | 221 |  | 63.33 | 107.67 | 22483 | 0.00153 | 0.00138 | 0.153 | 0.0138 | H | neg | neg
143 | PAMF Pre-COVID | Negative | 222 |  | -0.33 | 146.83 | 18249.17 | -0.0016 | 0.00384 | -0.16 | 0.0384 | H | neg | neg
144 | PAMF Pre-COVID | Negative | 223 |  | 2.33 | 102.5 | 20074.33 | -0.00132 | 0.00129 | -0.132 | 0.0129 | H | neg | neg
145 | PAMF Pre-COVID | Negative | 224 |  | -45.67 | 101 | 21384.67 | -0.00228 | 0.00114 | -0.228 | 0.0114 | H | neg | neg
146 | PAMF Pre-COVID | Negative | 225 |  | -32.67 | 66.67 | 19855.83 | -0.0018 | -0.0005 | -0.18 | -0.005 | H | neg | neg
147 | PAMF Pre-COVID | Negative | 226 |  | -48.5 | 95.67 | 19735 | -0.00261 | 0.00096 | -0.261 | 0.0096 | H | neg | neg
148 | PAMF Pre-COVID | Negative | 227 |  | -9.67 | 100 | 19438.67 | -0.00065 | 0.0012 | -0.065 | 0.012 | H | neg | neg
149 | PAMF Pre-COVID | Negative | 228 |  | -103.83 | 110 | 18198 | -0.00587 | 0.00183 | -0.587 | 0.0183 | H | neg | neg
150 | PAMF Pre-COVID | Negative | 229 |  | 2.33 | 298.33 | 18780.83 | -4e-05 | 0.0118 | -0.004 | 0.118 | H | neg | neg
151 | PAMF Pre-COVID | Negative | 230 |  | 9.67 | 2 | 19837.83 | 0.00034 | -0.00376 | 0.034 | -0.0376 | H | neg | neg
152 | PAMF Pre-COVID | Negative | 231 |  | -0.5 | 91.67 | 18954.17 | -0.00018 | 0.00079 | -0.018 | 0.0079 | H | neg | neg
153 | PAMF Pre-COVID | Negative | 232 |  | -3 | 102 | 17700.33 | -0.00034 | 0.00143 | -0.034 | 0.0143 | H | neg | neg
154 | PAMF Pre-COVID | Negative | 233 |  | 28.33 | 142 | 20515.5 | 0.00123 | 0.00318 | 0.123 | 0.0318 | H | neg | neg
155 | PAMF Pre-COVID | Negative | 234 |  | -0.67 | 100.5 | 17600.17 | -0.00021 | 0.00135 | -0.021 | 0.0135 | H | neg | neg
156 | PAMF Pre-COVID | Negative | 235 |  | 22 | 85.67 | 19212.17 | 0.00099 | 0.00047 | 0.099 | 0.0047 | H | neg | neg
157 | PAMF Pre-COVID | Negative | 236 |  | -32 | 96 | 18721.83 | -0.00172 | 0.00248 | -0.172 | 0.0248 | H | neg | neg
158 | PAMF Pre-COVID | Negative | 237 |  | 14.67 | 107.33 | 16448.67 | 0.00088 | 0.00352 | 0.088 | 0.0352 | H | neg | neg
159 | PAMF Pre-COVID | Negative | 238 |  | -104.67 | 86 | 13723.33 | -0.00764 | 0.00266 | -0.764 | 0.0266 | H | neg | neg
160 | PAMF Pre-COVID | Negative | 239 |  | -1 | 132.83 | 13820.5 | -8e-05 | 0.00603 | -0.008 | 0.0603 | H | neg | neg
161 | PAMF Pre-COVID | Negative | 240 |  | -6 | 102.83 | 15764.33 | -0.00039 | 0.00338 | -0.039 | 0.0338 | H | neg | neg
162 | PAMF Pre-COVID | Negative | 241 |  | -34 | 91.33 | 14834 | -0.0023 | 0.00282 | -0.23 | 0.0282 | H | neg | neg
163 | PAMF Pre-COVID | Negative | 242 |  | -8.33 | 89.17 | 14473.17 | -0.00059 | 0.00274 | -0.059 | 0.0274 | H | neg | neg
164 | PAMF Pre-COVID | Negative | 243 |  | -5.17 | 88 | 16581.83 | -0.00032 | 0.00232 | -0.032 | 0.0232 | H | neg | neg
165 | PAMF Pre-COVID | Negative | 244 |  | 1.17 | 158.67 | 15070.17 | 7e-05 | 0.00724 | 0.007 | 0.0724 | H | neg | neg
166 | PAMF Pre-COVID | Negative | 245 |  | -405 | 104 | 21562.83 | -0.01879 | 0.00253 | -1.879 | 0.0253 | H | neg | neg
167 | PAMF Pre-COVID | Negative | 247 |  | 168 | 118.3333 | 12739.8333 | 0.013056 | 0.0034799 | 1.3056 | 0.034799 | H | neg | neg
168 | PAMF Pre-COVID | Negative | 248 |  | -109.5 | 3 | 18952.33 | -0.00592 | -0.00153 | -0.592 | -0.0153 | H | neg | neg
169 | PAMF Pre-COVID | Negative | 249 |  | -97 | 206.33 | 17672.83 | -0.00564 | 0.00986 | -0.564 | 0.0986 | H | neg | neg
170 | PAMF Pre-COVID | Negative | 250 |  | -107 | 85.67 | 15799.33 | -0.00694 | 0.0034 | -0.694 | 0.034 | H | neg | neg
171 | PAMF Pre-COVID | Negative | 251 |  | -113 | 88.67 | 17219.5 | -0.00672 | 0.00329 | -0.672 | 0.0329 | H | neg | neg
172 | PAMF Pre-COVID | Negative | 252 |  | 0.33 | 68.67 | 15872.67 | -0.00015 | 0.00231 | -0.015 | 0.0231 | H | neg | neg
173 | PAMF Pre-COVID | Negative | 253 |  | -3.67 | 97 | 17081.5 | -0.00037 | 0.00381 | -0.037 | 0.0381 | H | neg | neg
174 | PAMF Pre-COVID | Negative | 254 |  | -0.67 | 99 | 19401.33 | -0.00017 | 0.00345 | -0.017 | 0.0345 | H | neg | neg
175 | PAMF Pre-COVID | Negative | 255 |  | 4.83 | 102.5 | 19562.5 | 0.00011 | 0.0036 | 0.011 | 0.036 | H | neg | neg
176 | PAMF Pre-COVID | Negative | 256 |  | -1.67 | 85.83 | 17179.33 | -0.00025 | 0.00313 | -0.025 | 0.0313 | H | neg | neg
177 | PAMF Pre-COVID | Negative | 257 |  | 29.33 | 101.33 | 21580.67 | 0.00124 | 0.00321 | 0.124 | 0.0321 | H | neg | neg
178 | PAMF Pre-COVID | Negative | 258 |  | -2 | 92.33 | 17783.67 | -0.00026 | 0.00339 | -0.026 | 0.0339 | H | neg | neg
179 | PAMF Pre-COVID | Negative | 259 |  | 62.83 | 107.33 | 19244.17 | 0.00313 | 0.00391 | 0.313 | 0.0391 | H | neg | neg
180 | PAMF Pre-COVID | Negative | 260 |  | -9.5 | 100.17 | 18253.83 | -0.00076 | 0.00151 | -0.076 | 0.0151 | H | neg | neg
181 | PAMF Pre-COVID | Negative | 261 |  | -65.33 | 47.33 | 15633.17 | -0.00446 | -0.00162 | -0.446 | -0.0162 | H | neg | neg
182 | PAMF Pre-COVID | Negative | 262 |  | -39.33 | 67.17 | 16181.33 | -0.0027 | -0.00034 | -0.27 | -0.0034 | H | neg | neg
183 | PAMF Pre-COVID | Negative | 263 |  | 1.83 | 101.67 | 16241.83 | -0.00015 | 0.00179 | -0.015 | 0.0179 | H | neg | neg
184 | PAMF Pre-COVID | Negative | 264 |  | -1.83 | 87.5 | 16657.17 | -0.00037 | 0.00089 | -0.037 | 0.0089 | H | neg | neg
185 | PAMF Pre-COVID | Negative | 265 |  | 66.33 | 108 | 17490.5 | 0.00354 | 0.00202 | 0.354 | 0.0202 | H | neg | neg
186 | PAMF Pre-COVID | Negative | 266 |  | 4 | 45.17 | 17403.33 | -2e-05 | -0.00158 | -0.002 | -0.0158 | H | neg | neg
187 | PAMF Pre-COVID | Negative | 267 |  | 39.67 | 106.67 | 20756.33 | 0.0017 | 0.00164 | 0.17 | 0.0164 | H | neg | neg
188 | PAMF Pre-COVID | Negative | 268 |  | 107 | 64.33 | 18670.33 | 0.0055 | -0.00045 | 0.55 | -0.0045 | H | neg | neg
189 | PAMF Pre-COVID | Negative | 269 |  | 3.83 | 93 | 22080.17 | -2e-05 | 0.00092 | -0.002 | 0.0092 | H | neg | neg
190 | PAMF Pre-COVID | Negative | 270 |  | 67.67 | 100.67 | 17024.5 | 0.00372 | 0.00164 | 0.372 | 0.0164 | H | neg | neg
191 | PAMF Pre-COVID | Negative | 271 |  | 64.67 | 110.5 | 18109 | 0.00333 | 0.00209 | 0.333 | 0.0209 | H | neg | neg
192 | PAMF Pre-COVID | Negative | 272 |  | -103.33 | 13.33 | 16386.5 | -0.00649 | -0.00393 | -0.649 | -0.0393 | H | neg | neg
193 | PAMF Pre-COVID | Negative | 273 |  | -58 | 13.67 | 17368.17 | -0.00351 | -0.00368 | -0.351 | -0.0368 | H | neg | neg
194 | PAMF Pre-COVID | Negative | 274 |  | 6.33 | 94.17 | 16833 | 0.0002 | 0.00098 | 0.02 | 0.0098 | H | neg | neg
195 | PAMF Pre-COVID | Negative | 275 |  | -2.33 | 96.83 | 17430.5 | -0.00031 | 0.0011 | -0.031 | 0.011 | H | neg | neg
196 | PAMF Pre-COVID | Negative | 276 |  | -2.67 | 110.17 | 17963 | -0.00032 | 0.00181 | -0.032 | 0.0181 | H | neg | neg
197 | PAMF Pre-COVID | Negative | 277 |  | 3.33 | 102.5 | 18504.67 | 2e-05 | 0.00134 | 0.002 | 0.0134 | H | neg | neg
198 | PAMF Pre-COVID | Negative | 278 |  | -2.83 | 13 | 20290.17 | -0.00029 | -0.00319 | -0.029 | -0.0319 | H | neg | neg
199 | PAMF Pre-COVID | Negative | 279 |  | 59.33 | 72 | 20903.33 | 0.00269 | -0.00027 | 0.269 | -0.0027 | H | neg | neg
200 | PAMF Pre-COVID | Negative | 280 |  | 3.67 | 15 | 19366.33 | 3e-05 | -0.00324 | 0.003 | -0.0324 | H | neg | neg
201 | PAMF Pre-COVID | Negative | 281 |  | 56.33 | 399.17 | 21866.67 | 0.00244 | 0.0147 | 0.244 | 0.147 | H | neg | neg
202 | PAMF Pre-COVID | Negative | 282 |  | 43 | 107.17 | 18471.83 | 0.00217 | 0.0016 | 0.217 | 0.016 | H | neg | neg
203 | PAMF Pre-COVID | Negative | 283 |  | 33.33 | 78 | 21685 | 0.0014 | 2e-05 | 0.14 | 0.0002 | H | neg | neg
204 | PAMF Pre-COVID | Negative | 284 |  | -98 | 67.5 | 17306.67 | -0.0063 | -0.00065 | -0.63 | -0.0065 | H | neg | neg
205 | PAMF Pre-COVID | Negative | 285 |  | -3.17 | 9.67 | 15996.5 | -0.00089 | -0.00431 | -0.089 | -0.0431 | H | neg | neg
206 | PAMF Pre-COVID | Negative | 286 |  | -3.67 | 4.67 | 16829.17 | -0.00087 | -0.0044 | -0.087 | -0.044 | H | neg | neg
207 | PAMF Pre-COVID | Negative | 287 |  | 0.83 | 110.33 | 17600.67 | -0.00058 | 0.0018 | -0.058 | 0.018 | H | neg | neg
208 | PAMF Pre-COVID | Negative | 288 |  | 2 | 43 | 18249.17 | -0.00049 | -0.00195 | -0.049 | -0.0195 | H | neg | neg
209 | PAMF Pre-COVID | Negative | 289 |  | 9.67 | 138.5 | 18586.33 | -7e-05 | 0.00322 | -0.007 | 0.0322 | H | neg | neg
210 | PAMF Pre-COVID | Negative | 290 |  | -2 | 110.67 | 19234.33 | -0.00068 | 0.00166 | -0.068 | 0.0166 | H | neg | neg
211 | PAMF Pre-COVID | Negative | 291 |  | -2 | 81.67 | 22343.17 | -0.00058 | 0.00013 | -0.058 | 0.0013 | H | neg | neg
212 | PAMF Pre-COVID | Negative | 292 |  | -0.67 | 78.33 | 19945.67 | -0.00058 | -2e-05 | -0.058 | -0.0002 | H | neg | neg
213 | PAMF Pre-COVID | Negative | 293 |  | -0.83 | 82.5 | 23204.33 | -0.00051 | 0.00017 | -0.051 | 0.0017 | H | neg | neg
214 | PAMF Pre-COVID | Negative | 294 |  | -33.5 | 102 | 17095.83 | -0.0026 | 0.00136 | -0.26 | 0.0136 | H | neg | neg
215 | PAMF Pre-COVID | Negative | 295 |  | 5 | 121.5 | 24044.33 | -0.00025 | 0.00178 | -0.025 | 0.0178 | H | neg | neg
216 | PAMF Pre-COVID | Negative | 382 |  | 184.5 | 270.3333 | 13430 | 0.013614 | 0.014619 | 1.3614 | 0.14619 | H | neg | neg
217 | PAMF Pre-COVID | Negative | 384 |  | 67.6667 | 230.3333 | 13348.8333 | 0.0049443 | 0.011711 | 0.49443 | 0.11711 | H | neg | neg
218 | PAMF Pre-COVID | Negative | 390 |  | 0.33333 | 92.3333 | 22338 | -0.0030814 | 0.0011042 | -0.30814 | 0.011042 | H | neg | neg
219 | PAMF Pre-COVID | Negative | 391 |  | 125 | 112.3333 | 13523 | 0.00668 | 0.004289 | 0.668 | 0.04289 | H | neg | neg
220 | PAMF Pre-COVID | Negative | 392 |  | 137.5 | 78 | 13608.1667 | 0.0075567 | 0.0017392 | 0.75567 | 0.017392 | H | neg | neg
221 | PAMF Pre-COVID | Negative | 393 |  | 70.6667 | 106.6667 | 13685.6667 | 0.0026305 | 0.003824 | 0.26305 | 0.03824 | H | neg | neg
222 | PAMF Pre-COVID | Negative | 394 |  | 73.3333 | 401.1667 | 11863.5 | 0.0032593 | 0.029235 | 0.32593 | 0.29235 | H | neg | neg
223 | PAMF Pre-COVID | Negative | 395 |  | 205.3333 | 104.6667 | 13438.6667 | 0.0127 | 0.0037454 | 1.27 | 0.037454 | H | neg | neg
224 | PAMF Pre-COVID | Negative | 396 |  | 90.3333 | 104.6667 | 12232.6667 | 0.0045507 | 0.0041147 | 0.45507 | 0.041147 | H | neg | neg
225 | PAMF Pre-COVID | Negative | 397 |  | 97.5 | 68.3333 | 13940.8333 | 0.0045071 | 0.0010042 | 0.45071 | 0.010042 | H | neg | neg
226 | PAMF Pre-COVID | Negative | 398 |  | 155.1667 | 110.3333 | 12392.6667 | 0.0097235 | 0.0045188 | 0.97235 | 0.045188 | H | neg | neg
227 | PAMF Pre-COVID | Negative | 399 |  | 73.5 | 134 | 12691.6667 | 0.0030598 | 0.0062771 | 0.30598 | 0.062771 | H | neg | neg
228 | PAMF Pre-COVID | Negative | 400 |  | -43.1667 | 103 | 17002.6667 | -0.0022644 | 0.0051953 | -0.22644 | 0.051953 | H | neg | neg
229 | PAMF Pre-COVID | Negative | 401 |  | -202.3333 | 86 | 15399.8333 | -0.012836 | 0.0046321 | -1.2836 | 0.046321 | H | neg | neg
230 | PAMF Pre-COVID | Negative | 402 |  | -8.6667 | 58.3333 | 15538.8333 | -0.00025742 | 0.0028102 | -0.025742 | 0.028102 | H | neg | neg
231 | PAMF Pre-COVID | Negative | 403 |  | 1.6667 | 85.6667 | 13433.8333 | 0.00047145 | 0.0052852 | 0.047145 | 0.052852 | H | neg | neg
232 | PAMF Pre-COVID | Negative | 404 |  | -13.6667 | 159 | 14607 | -0.00061614 | 0.0098811 | -0.061614 | 0.098811 | H | neg | neg
233 | PAMF Pre-COVID | Negative | 405 |  | 0.83333 | 109 | 14789.5 | 0.00037189 | 0.0063784 | 0.037189 | 0.063784 | H | neg | neg
234 | PAMF Pre-COVID | Negative | 406 |  | -4.6667 | 97.3333 | 12936.3333 | 0 | 0.0063903 | 0 | 0.063903 | H | neg | neg
235 | PAMF Pre-COVID | Negative | 407 |  | 58 | 89.3333 | 13796.3333 | 0.0045423 | 0.0054121 | 0.45423 | 0.054121 | H | neg | neg
236 | PAMF Pre-COVID | Negative | 408 |  | 39.3333 | 91 | 12456.1667 | 0.0035324 | 0.0061282 | 0.35324 | 0.061282 | H | neg | neg
237 | PAMF Pre-COVID | Negative | 409 |  | 10.1667 | 80.6667 | 12993 | 0.0011416 | 0.0050797 | 0.11416 | 0.050797 | H | neg | neg
238 | PAMF Pre-COVID | Negative | 410 |  | 0.66667 | 92.3333 | 12688 | 0.00042034 | 0.0061213 | 0.042034 | 0.061213 | H | neg | neg
239 | PAMF Pre-COVID | Negative | 411 |  | 63.5 | 90.8333 | 13296.3333 | 0.0051267 | 0.0057284 | 0.51267 | 0.057284 | H | neg | neg
240 | PAMF Pre-COVID | Negative | 412 |  | -105.3333 | 83.6667 | 17555.5 | -0.0063323 | 0.0006076 | -0.63323 | 0.006076 | H | neg | neg
241 | PAMF Pre-COVID | Negative | 413 |  | -34.6667 | 12.8333 | 18262 | -0.0022177 | -0.0032946 | -0.22177 | -0.032946 | H | neg | neg
242 | PAMF Pre-COVID | Negative | 414 |  | 5.3333 | 83.3333 | 16007.1667 | -3.1236e-05 | 0.00064554 | -0.0031236 | 0.0064554 | H | neg | neg
243 | PAMF Pre-COVID | Negative | 415 |  | -5.6667 | 76 | 16834 | -0.00068314 | 0.00017821 | -0.068314 | 0.0017821 | H | neg | neg
244 | PAMF Pre-COVID | Negative | 416 |  | -35.3333 | 40.6667 | 16613.8333 | -0.0024779 | -0.0019462 | -0.24779 | -0.019462 | H | neg | neg
245 | PAMF Pre-COVID | Negative | 417 |  | 0 | 184.3333 | 16928.6667 | -0.00034458 | 0.0065766 | -0.034458 | 0.065766 | H | neg | neg
246 | PAMF Pre-COVID | Negative | 418 |  | 11 | 90.6667 | 14634.3333 | 0.00035305 | 0.0012072 | 0.035305 | 0.012072 | H | neg | neg
247 | PAMF Pre-COVID | Negative | 419 |  | -0.16667 | 1747.1667 | 15465.1667 | -0.00038797 | 0.10825 | -0.038797 | 1.0825 | H | neg | neg
248 | PAMF Pre-COVID | Negative | 420 |  | -6.6667 | 91.3333 | 13739 | -0.00090982 | 0.0013344 | -0.090982 | 0.013344 | H | neg | neg
249 | PAMF Pre-COVID | Negative | 421 |  | -7.8333 | 119.3333 | 16254.1667 | -0.00084081 | 0.0028506 | -0.084081 | 0.028506 | H | neg | neg
250 | PAMF Pre-COVID | Negative | 422 |  | 1.8333 | 57.3333 | 13990.1667 | -0.00028592 | -0.0011198 | -0.028592 | -0.011198 | H | neg | neg
251 | PAMF Pre-COVID | Negative | 423 |  | 89.6667 | 126 | 16117.8333 | 0.0052013 | 0.0032883 | 0.52013 | 0.032883 | H | neg | neg
252 | PAMF Pre-COVID | Negative | 424 |  | -69.6667 | 121 | 18390.6667 | -0.0054738 | 0.0025647 | -0.54738 | 0.025647 | H | neg | neg
253 | PAMF Pre-COVID | Negative | 425 |  | -32.3333 | 101 | 18836.1667 | -0.0033623 | 0.0014423 | -0.33623 | 0.014423 | H | neg | neg
254 | PAMF Pre-COVID | Negative | 426 |  | 0 | 81.3333 | 17348.1667 | -0.0017869 | 0.00043232 | -0.17869 | 0.0043232 | H | neg | neg
255 | PAMF Pre-COVID | Negative | 427 |  | -0.5 | 156 | 18270.8333 | -0.0017241 | 0.0044971 | -0.17241 | 0.044971 | H | neg | neg
256 | PAMF Pre-COVID | Negative | 428 |  | -6.1667 | 101.6667 | 19137.8333 | -0.0019421 | 0.0014544 | -0.19421 | 0.014544 | H | neg | neg
257 | PAMF Pre-COVID | Negative | 429 |  | -0.83333 | 88 | 18800.3333 | -0.0016932 | 0.00075353 | -0.16932 | 0.0075353 | H | neg | neg
258 | PAMF Pre-COVID | Negative | 430 |  | 1 | 157.8333 | 17776.6667 | -0.0016876 | 0.0047253 | -0.16876 | 0.047253 | H | neg | neg
259 | PAMF Pre-COVID | Negative | 431 |  | 103.6667 | 196.8333 | 16578.5 | 0.0043832 | 0.0074192 | 0.43832 | 0.074192 | H | neg | neg
260 | PAMF Pre-COVID | Negative | 432 |  | 109.3333 | 96.5 | 16218.3333 | 0.0048299 | 0.0013976 | 0.48299 | 0.013976 | H | neg | neg
261 | PAMF Pre-COVID | Negative | 433 |  | 100 | 17.5 | 17573.1667 | 0.0039264 | -0.0032056 | 0.39264 | -0.032056 | H | neg | neg
262 | PAMF Pre-COVID | Negative | 434 |  | 279.5 | 3180.5 | 13540.3333 | 0.020519 | 0.22943 | 2.0519 | 2.2943 | H | pos | pos
263 | PAMF Pre-COVID | Negative | 435 |  | 4 | 64 | 16584.5 | -0.001628 | -0.00059292 | -0.1628 | -0.0059292 | H | neg | neg
264 | PAMF Pre-COVID | Negative | 436 |  | 104 | 359.6667 | 12917.1667 | 0.0079223 | 0.022115 | 0.79223 | 0.22115 | H | neg | neg
265 | PAMF Pre-COVID | Negative | 437 |  | -9 | 86.6667 | 15454.8333 | -0.001251 | 0.00069018 | -0.1251 | 0.0069018 | H | neg | neg
266 | PAMF Pre-COVID | Negative | 438 |  | -104 | 84.3333 | 15737.8333 | -0.0072649 | 0.00052951 | -0.72649 | 0.0052951 | H | neg | neg
267 | PAMF Pre-COVID | Negative | 439 |  | -82.3333 | 40.8333 | 16517 | -0.0056104 | -0.0021291 | -0.56104 | -0.021291 | H | neg | neg
268 | PAMF Pre-COVID | Negative | 440 |  | -6.1667 | 102.3333 | 17890.6667 | -0.00092227 | 0.0014719 | -0.092227 | 0.014719 | H | neg | neg
269 | PAMF Pre-COVID | Negative | 441 |  | -34.1667 | 80.6667 | 19367.6667 | -0.0022976 | 0.00024095 | -0.22976 | 0.0024095 | H | neg | neg
270 | PAMF Pre-COVID | Negative | 442 |  | -6 | 101.3333 | 18910.8333 | -0.0008637 | 0.0013396 | -0.08637 | 0.013396 | H | neg | neg
271 | PAMF Pre-COVID | Negative | 443 |  | 12 | 97.5 | 18771.1667 | 8.8789e-05 | 0.0011454 | 0.0088789 | 0.011454 | H | neg | neg
272 | PAMF Pre-COVID | Negative | 444 |  | -2.6667 | 90.3333 | 19214.1667 | -0.00067658 | 0.00074598 | -0.067658 | 0.0074598 | H | neg | neg
273 | PAMF Pre-COVID | Negative | 445 |  | 12.1667 | 185.8333 | 20510.3333 | 8.9386e-05 | 0.005355 | 0.0089386 | 0.05355 | H | neg | neg
274 | PAMF Pre-COVID | Negative | 446 |  | 102.6667 | 89.3333 | 19281.1667 | 0.0047888 | 0.00069152 | 0.47888 | 0.0069152 | H | neg | neg
275 | PAMF Pre-COVID | Negative | 447 |  | -0.33333 | 125.6667 | 19724.1667 | -0.00054079 | 0.0025181 | -0.054079 | 0.025181 | H | neg | neg
276 | PAMF Pre-COVID | Negative | 448 |  | 6.1667 | 60.6667 | 27832.3333 | -0.0022695 | -0.00037127 | -0.22695 | -0.0037127 | H | neg | neg
277 | PAMF Pre-COVID | Negative | 449 |  | 90.3333 | 89.1667 | 26702.6667 | 0.00078644 | 0.00068033 | 0.078644 | 0.0068033 | H | neg | neg
278 | PAMF Pre-COVID | Negative | 450 |  | -33.8333 | 68.5 | 24334.6667 | -0.0042395 | -0.00010273 | -0.42395 | -0.0010273 | H | neg | neg
279 | PAMF Pre-COVID | Negative | 451 |  | -2 | 15 | 22840.3333 | -0.0031231 | -0.0024518 | -0.31231 | -0.024518 | H | neg | neg
280 | PAMF Pre-COVID | Negative | 452 |  | 139.1667 | 67 | 23203.1667 | 0.0030096 | -0.00017239 | 0.30096 | -0.0017239 | H | neg | neg
281 | PAMF Pre-COVID | Negative | 453 |  | 58.3333 | 90 | 21597.5 | -0.00050932 | 0.00087973 | -0.050932 | 0.0087973 | H | neg | neg
282 | PAMF Pre-COVID | Negative | 454 |  | -1.3333 | 177.3333 | 22008.6667 | -0.0032109 | 0.0048314 | -0.32109 | 0.048314 | H | neg | neg
283 | PAMF Pre-COVID | Negative | 455 |  | 94.3333 | 91.5 | 21536.8333 | 0.0011608 | 0.00095186 | 0.11608 | 0.0095186 | H | neg | neg
284 | PAMF Pre-COVID | Negative | 456 |  | 3 | 93 | 23460.6667 | -0.0028274 | 0.00093774 | -0.28274 | 0.0093774 | H | neg | neg
285 | PAMF Pre-COVID | Negative | 457 |  | 25.8333 | 91.5 | 21935.6667 | -0.0019831 | 0.00093455 | -0.19831 | 0.0093455 | H | neg | neg
286 | PAMF Pre-COVID | Negative | 458 |  | 107.1667 | 275.5 | 21993.3333 | 0.0017202 | 0.0092983 | 0.17202 | 0.092983 | H | neg | neg
287 | PAMF Pre-COVID | Negative | 459 |  | 105.3333 | 197.5 | 22242 | 0.0016186 | 0.0056874 | 0.16186 | 0.056874 | H | neg | neg
288 | PAMF Pre-COVID | Negative | 460 |  | 89.6667 | 93.6667 | 26604.3333 | 0.0022678 | 0.00055129 | 0.22678 | 0.0055129 | H | neg | neg
289 | PAMF Pre-COVID | Negative | 461 |  | 103.6667 | 103.3333 | 26880 | 0.0027654 | 0.00090526 | 0.27654 | 0.0090526 | H | neg | neg
290 | PAMF Pre-COVID | Negative | 462 |  | 0 | 93.3333 | 24194.8333 | -0.0012124 | 0.00059241 | -0.12124 | 0.0059241 | H | neg | neg
291 | PAMF Pre-COVID | Negative | 463 |  | -2.3333 | 95.3333 | 23432.6667 | -0.0013514 | 0.00069703 | -0.13514 | 0.0069703 | H | neg | neg
292 | PAMF Pre-COVID | Negative | 464 |  | 97 | 97.6667 | 23576.6667 | 0.0028701 | 0.00079174 | 0.28701 | 0.0079174 | H | neg | neg
293 | PAMF Pre-COVID | Negative | 465 |  | 94.3333 | 155.3333 | 21593 | 0.0030102 | 0.0035351 | 0.30102 | 0.035351 | H | neg | neg
294 | PAMF Pre-COVID | Negative | 466 |  | 102.6667 | 219.6667 | 23017 | 0.0031861 | 0.0061114 | 0.31861 | 0.061114 | H | neg | neg
295 | PAMF Pre-COVID | Negative | 467 |  | 72 | 67.3333 | 18886.8333 | 0.0022591 | -0.00061771 | 0.22591 | -0.0061771 | H | neg | neg
296 | PAMF Pre-COVID | Negative | 468 |  | 101.3333 | 156.5 | 22710.6667 | 0.0031703 | 0.0034125 | 0.31703 | 0.034125 | H | neg | neg
297 | PAMF Pre-COVID | Negative | 469 |  | 30 | 92.3333 | 19578.3333 | 3.4051e-05 | 0.00068102 | 0.0034051 | 0.0068102 | H | neg | neg
298 | PAMF Pre-COVID | Negative | 470 |  | 117 | 110.1667 | 20032.8333 | 0.0043761 | 0.0015558 | 0.43761 | 0.015558 | H | neg | neg
299 | PAMF Pre-COVID | Negative | 471 |  | 93 | 93.3333 | 20123.3333 | 0.0031638 | 0.00071227 | 0.31638 | 0.0071227 | H | neg | neg
300 | PAMF Pre-COVID | Negative | 472 |  | -280 | 101.6667 | 27230.8333 | -0.013906 | 0.00078343 | -1.3906 | 0.0078343 | H | neg | neg
301 | PAMF Pre-COVID | Negative | 473 |  | 56.6667 | 95.3333 | 29140.1667 | -0.0014413 | 0.00051475 | -0.14413 | 0.0051475 | H | neg | neg
302 | PAMF Pre-COVID | Negative | 474 |  | -4.5 | 94 | 26540 | -0.0038872 | 0.00051495 | -0.38872 | 0.0051495 | H | neg | neg
303 | PAMF Pre-COVID | Negative | 475 |  | 4 | 33.6667 | 25150.3333 | -0.003764 | -0.0018555 | -0.3764 | -0.018555 | H | neg | neg
304 | PAMF Pre-COVID | Negative | 476 |  | 37.5 | 204.6667 | 26171 | -0.0023372 | 0.0047508 | -0.23372 | 0.047508 | H | neg | neg
305 | PAMF Pre-COVID | Negative | 477 |  | 2 | 93.3333 | 24194 | -0.0039955 | 0.00053732 | -0.39955 | 0.0053732 | H | neg | neg
306 | PAMF Pre-COVID | Negative | 478 |  | 95.3333 | 191.6667 | 26561.5 | -0.00012549 | 0.0041915 | -0.012549 | 0.041915 | H | neg | neg
307 | PAMF Pre-COVID | Negative | 479 |  | 69.3333 | 105.1667 | 24794 | -0.0011831 | 0.0010016 | -0.11831 | 0.010016 | H | neg | neg
308 | PAMF Pre-COVID | Negative | 480 |  | 94.8333 | 108.1667 | 25701.1667 | -0.00014915 | 0.001083 | -0.014915 | 0.01083 | H | neg | neg
309 | PAMF Pre-COVID | Negative | 481 |  | 68 | 92 | 25728.3333 | -0.0011919 | 0.00045346 | -0.11919 | 0.0045346 | H | neg | neg
310 | PAMF Pre-COVID | Negative | 482 |  | 97.6667 | 97 | 22704 | -4.4045e-05 | 0.00073409 | -0.0044045 | 0.0073409 | H | neg | neg
311 | PAMF Pre-COVID | Negative | 483 |  | 8.3333 | 105.6667 | 23337.8333 | -0.0038707 | 0.0010855 | -0.38707 | 0.010855 | H | neg | neg
312 | PAMF Pre-COVID | Negative | 484 |  | -36 | 87.6667 | 31071.5 | -0.0012123 | 0.00030038 | -0.12123 | 0.0030038 | H | neg | neg
313 | PAMF Pre-COVID | Negative | 485 |  | -0.83333 | 90.8333 | 33278.3333 | -7.5124e-05 | 0.00037562 | -0.0075124 | 0.0037562 | H | neg | neg
314 | PAMF Pre-COVID | Negative | 486 |  | -7 | 85.6667 | 27169.1667 | -0.00031899 | 0.00026991 | -0.031899 | 0.0026991 | H | neg | neg
315 | PAMF Pre-COVID | Negative | 487 |  | 104.6667 | 291 | 13170.3333 | 0.0078206 | 0.016476 | 0.78206 | 0.16476 | H | neg | neg
316 | PAMF Pre-COVID | Negative | 488 |  | 5 | 177.1667 | 27447 | 0.00012145 | 0.0036009 | 0.012145 | 0.036009 | H | neg | neg
317 | PAMF Pre-COVID | Negative | 489 |  | 31 | 104.5 | 29301.5 | 0.0010011 | 0.00089301 | 0.10011 | 0.0089301 | H | neg | neg
318 | PAMF Pre-COVID | Negative | 490 |  | -1.3333 | 109.6667 | 27390.6667 | -0.00010953 | 0.0011439 | -0.010953 | 0.011439 | H | neg | neg
319 | PAMF Pre-COVID | Negative | 491 |  | 27.3333 | 97.3333 | 27159.6667 | 0.00094503 | 0.00069957 | 0.094503 | 0.0069957 | H | neg | neg
320 | PAMF Pre-COVID | Negative | 492 |  | 59.5 | 96 | 26178.8333 | 0.0022092 | 0.00067485 | 0.22092 | 0.0067485 | H | neg | neg
321 | PAMF Pre-COVID | Negative | 493 |  | 58.1667 | 91.3333 | 27881.3333 | 0.0020264 | 0.00046626 | 0.20264 | 0.0046626 | H | neg | neg
322 | PAMF Pre-COVID | Negative | 494 |  | 66.6667 | 101.3333 | 24358.1667 | 0.0026685 | 0.00094424 | 0.26685 | 0.0094424 | H | neg | neg
323 | PAMF Pre-COVID | Negative | 495 |  | 34 | 90.3333 | 25766.5 | 0.0012549 | 0.00046572 | 0.12549 | 0.0046572 | H | neg | neg
324 | PAMF Pre-COVID | Negative | 496 |  | -60 | 69 | 21773.6667 | -0.0028322 | 0.0024035 | -0.28322 | 0.024035 | H | neg | neg
325 | PAMF Pre-COVID | Negative | 497 |  | -5.1667 | 40.6667 | 19059.5 | -0.00035853 | 0.0012592 | -0.035853 | 0.012592 | H | neg | neg
326 | PAMF Pre-COVID | Negative | 498 |  | -40 | 94.1667 | 18153.1667 | -0.0022953 | 0.0042692 | -0.22953 | 0.042692 | H | neg | neg
327 | PAMF Pre-COVID | Negative | 499 |  | 1 | 60.3333 | 16761.8333 | -3.9773e-05 | 0.0026051 | -0.0039773 | 0.026051 | H | neg | neg
328 | PAMF Pre-COVID | Negative | 500 |  | -2.1667 | 69.3333 | 18523.3333 | -0.00020695 | 0.0028433 | -0.020695 | 0.028433 | H | neg | neg
329 | PAMF Pre-COVID | Negative | 501 |  | 1.3333 | 83.3333 | 16648.3333 | -2.0022e-05 | 0.0040044 | -0.0020022 | 0.040044 | H | neg | neg
330 | PAMF Pre-COVID | Negative | 502 |  | 67.6667 | 88.5 | 17442.6667 | 0.0037838 | 0.0041183 | 0.37838 | 0.041183 | H | neg | neg
331 | PAMF Pre-COVID | Negative | 503 |  | 0.83333 | 91 | 16846.8333 | -4.9465e-05 | 0.0044123 | -0.0049465 | 0.044123 | H | neg | neg
332 | PAMF Pre-COVID | Negative | 504 |  | 99.3333 | 84.3333 | 16312 | 0.0059874 | 0.0041483 | 0.59874 | 0.041483 | H | neg | neg
333 | PAMF Pre-COVID | Negative | 505 |  | 95.6667 | 103 | 18434.3333 | 0.0050992 | 0.0046833 | 0.50992 | 0.046833 | H | neg | neg
334 | PAMF Pre-COVID | Negative | 506 |  | -4.6667 | 108 | 17002 | -0.00037251 | 0.0053719 | -0.037251 | 0.053719 | H | neg | neg
335 | PAMF Pre-COVID | Negative | 507 |  | 68.6667 | 59.6667 | 16403.8333 | 0.0040844 | 0.0026213 | 0.40844 | 0.026213 | H | neg | neg
336 | PAMF Pre-COVID | Negative | 508 |  | -32 | 95.3333 | 23029.6667 | -0.0032567 | 0.00098424 | -0.32567 | 0.0098424 | H | neg | neg
337 | PAMF Pre-COVID | Negative | 509 |  | 3.5 | 88 | 22634.6667 | -0.0017451 | 0.00067743 | -0.17451 | 0.0067743 | H | neg | neg
338 | PAMF Pre-COVID | Negative | 510 |  | -34.5 | 16.1667 | 20495.3333 | -0.0037813 | -0.0027567 | -0.37813 | -0.027567 | H | neg | neg
339 | PAMF Pre-COVID | Negative | 511 |  | 3 | 99.3333 | 21285.8333 | -0.0018792 | 0.0012528 | -0.18792 | 0.012528 | H | neg | neg
340 | PAMF Pre-COVID | Negative | 512 |  | 100.6667 | 103 | 19458.6667 | 0.0029635 | 0.0015589 | 0.29635 | 0.015589 | H | neg | neg
341 | PAMF Pre-COVID | Negative | 513 |  | -3 | 86.3333 | 19924.5 | -0.0023087 | 0.00068592 | -0.23087 | 0.0068592 | H | neg | neg
342 | PAMF Pre-COVID | Negative | 514 |  | -66 | 67 | 19055.5 | -0.0057201 | -0.00029738 | -0.57201 | -0.0029738 | H | neg | neg
343 | PAMF Pre-COVID | Negative | 515 |  | 64.6667 | 90.3333 | 18173.5 | 0.0011922 | 0.00097211 | 0.11922 | 0.0097211 | H | neg | neg
344 | PAMF Pre-COVID | Negative | 516 |  | 7.6667 | 69 | 18646.5 | -0.0018949 | -0.00019664 | -0.18949 | -0.0019664 | H | neg | neg
345 | PAMF Pre-COVID | Negative | 517 |  | 72 | 96.3333 | 17635.1667 | 0.0016444 | 0.001342 | 0.16444 | 0.01342 | H | neg | neg
346 | PAMF Pre-COVID | Negative | 518 |  | -2.6667 | 121.5 | 17934.1667 | -0.0025464 | 0.0027229 | -0.25464 | 0.027229 | H | neg | neg
347 | PAMF Pre-COVID | Negative | 519 |  | 97.5 | 37.5 | 16707 | 0.0032621 | -0.0021049 | 0.32621 | -0.021049 | H | neg | neg
348 | PAMF Pre-COVID | Negative | 520 |  | 8.3333 | 62.5 | 27482.8333 | 0.00023045 | -0.00046696 | 0.023045 | -0.0046696 | H | neg | neg
349 | PAMF Pre-COVID | Negative | 521 |  | 105.3333 | 73.3333 | 27242.8333 | 0.003793 | -7.3414e-05 | 0.3793 | -0.00073414 | H | neg | neg
350 | PAMF Pre-COVID | Negative | 522 |  | -3.3333 | 15.1667 | 23286.3333 | -0.00022903 | -0.0025838 | -0.022903 | -0.025838 | H | neg | neg
351 | PAMF Pre-COVID | Negative | 523 |  | 98.3333 | 92.6667 | 22489.8333 | 0.0042834 | 0.00077072 | 0.42834 | 0.0077072 | H | neg | neg
352 | PAMF Pre-COVID | Negative | 524 |  | 2.6667 | 89.6667 | 22142.5 | 3.0108e-05 | 0.00064732 | 0.0030108 | 0.0064732 | H | neg | neg
353 | PAMF Pre-COVID | Negative | 525 |  | -0.33333 | 95 | 21737.1667 | -0.00010734 | 0.00090475 | -0.010734 | 0.0090475 | H | neg | neg
354 | PAMF Pre-COVID | Negative | 526 |  | 105 | 103.8333 | 20544.6667 | 0.0050135 | 0.0013872 | 0.50135 | 0.013872 | H | neg | neg
355 | PAMF Pre-COVID | Negative | 527 |  | 105.6667 | 88.6667 | 19984.6667 | 0.0051873 | 0.00066718 | 0.51873 | 0.0066718 | H | neg | neg
356 | PAMF Pre-COVID | Negative | 528 |  | 103.6667 | 103.3333 | 21841.1667 | 0.0046548 | 0.001282 | 0.46548 | 0.01282 | H | neg | neg
357 | PAMF Pre-COVID | Negative | 529 |  | 195.1667 | 69.1667 | 21978.5 | 0.0087889 | -0.00028058 | 0.87889 | -0.0028058 | H | neg | neg
358 | PAMF Pre-COVID | Negative | 530 |  | 2.6667 | 98.6667 | 21947 | 3.0376e-05 | 0.0010632 | 0.0030376 | 0.010632 | H | neg | neg
359 | PAMF Pre-COVID | Negative | 531 |  | 96 | 95 | 20707.6667 | 0.0045394 | 0.00094973 | 0.45394 | 0.0094973 | H | neg | neg
360 | PAMF Pre-COVID | Negative | 532 |  | 4.3333 | 102 | 30167 | -0.00020994 | 0.00083977 | -0.020994 | 0.0083977 | H | neg | neg
361 | PAMF Pre-COVID | Negative | 533 |  | 65 | 98.5 | 26593 | 0.0020431 | 0.00082102 | 0.20431 | 0.0082102 | H | neg | neg
362 | PAMF Pre-COVID | Negative | 534 |  | -68 | 94.8333 | 25630.1667 | -0.0030693 | 0.0007088 | -0.30693 | 0.007088 | H | neg | neg
363 | PAMF Pre-COVID | Negative | 535 |  | 72.3333 | 56.6667 | 26277.3333 | 0.0023468 | -0.00076111 | 0.23468 | -0.0076111 | H | neg | neg
364 | PAMF Pre-COVID | Negative | 536 |  | 191.6667 | 89 | 25460.8333 | 0.007109 | 0.0004844 | 0.7109 | 0.004844 | H | neg | neg
365 | PAMF Pre-COVID | Negative | 537 |  | -3 | 67 | 24574 | -0.00055614 | -0.00039337 | -0.055614 | -0.0039337 | H | neg | neg
366 | PAMF Pre-COVID | Negative | 538 |  | -11 | 352.8333 | 25328.8333 | -0.00085542 | 0.010903 | -0.085542 | 0.10903 | H | neg | neg
367 | PAMF Pre-COVID | Negative | 539 |  | 88 | 102.8333 | 22913.3333 | 0.003375 | 0.001142 | 0.3375 | 0.01142 | H | neg | neg
368 | PAMF Pre-COVID | Negative | 540 |  | 95.3333 | 130.1667 | 23963.1667 | 0.0035332 | 0.0022326 | 0.35332 | 0.022326 | H | neg | neg
369 | PAMF Pre-COVID | Negative | 541 |  | 75.1667 | 99 | 24881.1667 | 0.0025923 | 0.0008976 | 0.25923 | 0.008976 | H | neg | neg
370 | PAMF Pre-COVID | Negative | 542 |  | 96.3333 | 67 | 24297 | 0.0035258 | -0.00039785 | 0.35258 | -0.0039785 | H | neg | neg
371 | PAMF Pre-COVID | Negative | 543 |  | 97.6667 | 98.6667 | 24258.8333 | 0.0035863 | 0.00090689 | 0.35863 | 0.0090689 | H | neg | neg
372 | PAMF Pre-COVID | Negative | 544 |  | 35 | 79.6667 | 27111 | -0.0014754 | 0.00016598 | -0.14754 | 0.0016598 | H | neg | neg
373 | PAMF Pre-COVID | Negative | 545 |  | 124.3333 | 133.3333 | 24503.3333 | 0.0020133 | 0.0023738 | 0.20133 | 0.023738 | H | neg | neg
374 | PAMF Pre-COVID | Negative | 546 |  | -6.3333 | 88 | 25551 | -0.0031832 | 0.00050226 | -0.31832 | 0.0050226 | H | neg | neg
375 | PAMF Pre-COVID | Negative | 547 |  | 103.6667 | -1.3333 | 21601 | 0.0013271 | -0.0035415 | 0.13271 | -0.035415 | H | neg | neg
376 | PAMF Pre-COVID | Negative | 548 |  | 0.5 | 88.3333 | 21938.8333 | -0.0033958 | 0.00060015 | -0.33958 | 0.0060015 | H | neg | neg
377 | PAMF Pre-COVID | Negative | 549 |  | -2 | 105 | 22547.8333 | -0.003415 | 0.0013231 | -0.3415 | 0.013231 | H | neg | neg
378 | PAMF Pre-COVID | Negative | 550 |  | 40 | 94.3333 | 21491.3333 | -0.0016286 | 0.00089183 | -0.16286 | 0.0089183 | H | neg | neg
379 | PAMF Pre-COVID | Negative | 551 |  | 24.6667 | 100 | 20384 | -0.0024693 | 0.0012183 | -0.24693 | 0.012183 | H | neg | neg
380 | PAMF Pre-COVID | Negative | 552 |  | 90.6667 | 91 | 19873.6667 | 0.00078831 | 0.0007967 | 0.078831 | 0.007967 | H | neg | neg
381 | PAMF Pre-COVID | Negative | 553 |  | 104.6667 | 91.3333 | 20341 | 0.0014585 | 0.00079478 | 0.14585 | 0.0079478 | H | neg | neg
382 | PAMF Pre-COVID | Negative | 554 |  | 104.1667 | 108.8333 | 21483.6667 | 0.0013576 | 0.0015671 | 0.13576 | 0.015671 | H | neg | neg
383 | PAMF Pre-COVID | Negative | 555 |  | 99 | 44 | 18299.5 | 0.0013115 | -0.0017031 | 0.13115 | -0.017031 | H | neg | neg
384 | PAMF Pre-COVID | Negative | 556 |  | 5 | 40.6667 | 24075.8333 | -0.0023675 | -0.0013984 | -0.23675 | -0.013984 | H | neg | neg
385 | PAMF Pre-COVID | Negative | 557 |  | -0.66667 | 31.6667 | 22625.6667 | -0.0027697 | -0.0018858 | -0.27697 | -0.018858 | H | neg | neg
386 | PAMF Pre-COVID | Negative | 558 |  | 9 | 100.3333 | 21224.6667 | -0.0024971 | 0.001225 | -0.24971 | 0.01225 | H | neg | neg
387 | PAMF Pre-COVID | Negative | 559 |  | 72.5 | 251.3333 | 22985.3333 | 0.00045681 | 0.0077006 | 0.045681 | 0.077006 | H | neg | neg
388 | PAMF Pre-COVID | Negative | 560 |  | 83.5 | 102.1667 | 21236.1667 | 0.0010124 | 0.0013107 | 0.10124 | 0.013107 | H | neg | neg
389 | PAMF Pre-COVID | Negative | 561 |  | -0.33333 | 95.1667 | 20875.1667 | -0.002986 | 0.000998 | -0.2986 | 0.00998 | H | neg | neg
390 | PAMF Pre-COVID | Negative | 562 |  | 103 | 195.3333 | 22519.5 | 0.0018206 | 0.0053731 | 0.18206 | 0.053731 | H | neg | neg
391 | PAMF Pre-COVID | Negative | 563 |  | 96.8333 | 111.3333 | 22749.5 | 0.0015312 | 0.0016264 | 0.15312 | 0.016264 | H | neg | neg
392 | PAMF Pre-COVID | Negative | 564 |  | 2.1667 | 87 | 21964.1667 | -0.0027241 | 0.0005767 | -0.27241 | 0.005767 | H | neg | neg
393 | PAMF Pre-COVID | Negative | 565 |  | 136 | 101.3333 | 22404.6667 | 0.0033029 | 0.0012051 | 0.33029 | 0.012051 | H | neg | neg
394 | PAMF Pre-COVID | Negative | 566 |  | 104.6667 | 91 | 19201.5 | 0.002222 | 0.00086799 | 0.2222 | 0.0086799 | H | neg | neg
395 | PAMF Pre-COVID | Negative | 567 |  | 63.6667 | 94 | 21038.6667 | 7.9219e-05 | 0.00093479 | 0.0079219 | 0.0093479 | H | neg | neg
396 | PAMF Pre-COVID | Negative | 568 |  | 185.3333 | 10 | 25660.1667 | 0.0045076 | -0.0024811 | 0.45076 | -0.024811 | H | neg | neg
397 | PAMF Pre-COVID | Negative | 569 |  | -12.1667 | 114 | 27499.5 | -0.0029758 | 0.0014667 | -0.29758 | 0.014667 | H | neg | neg
398 | PAMF Pre-COVID | Negative | 570 |  | -3 | 64.3333 | 24153.1667 | -0.0030086 | -0.00038642 | -0.30086 | -0.0038642 | H | neg | neg
399 | PAMF Pre-COVID | Negative | 571 |  | 106.6667 | 86.6667 | 26860.8333 | 0.0013775 | 0.00048398 | 0.13775 | 0.0048398 | H | neg | neg
400 | PAMF Pre-COVID | Negative | 572 |  | 83.5 | 162.3333 | 23496.3333 | 0.00058874 | 0.0037736 | 0.058874 | 0.037736 | H | neg | neg
401 | PAMF Pre-COVID | Negative | 573 |  | -1.3333 | 108.3333 | 25122.1667 | -0.0028262 | 0.0013799 | -0.28262 | 0.013799 | H | neg | neg
402 | PAMF Pre-COVID | Negative | 574 |  | 41.6667 | 94.1667 | 24337.5 | -0.0011505 | 0.00084232 | -0.11505 | 0.0084232 | H | neg | neg
403 | PAMF Pre-COVID | Negative | 575 |  | 48.5 | 152.5 | 25661.5 | -0.00082484 | 0.003072 | -0.082484 | 0.03072 | H | neg | neg
404 | PAMF Pre-COVID | Negative | 576 |  | 101.1667 | 89 | 23302.5 | 0.0013518 | 0.00065801 | 0.13518 | 0.0065801 | H | neg | neg
405 | PAMF Pre-COVID | Negative | 577 |  | 4.3333 | 111 | 25319.8333 | -0.0025803 | 0.0014745 | -0.25803 | 0.014745 | H | neg | neg
406 | PAMF Pre-COVID | Negative | 578 |  | 106.1667 | 4.6667 | 21104 | 0.0017295 | -0.0032695 | 0.17295 | -0.032695 | H | neg | neg
407 | PAMF Pre-COVID | Negative | 579 |  | 102.5 | 102.3333 | 21519.1667 | 0.0015258 | 0.0013321 | 0.15258 | 0.013321 | H | neg | neg
408 | PAMF Pre-COVID | Negative | 580 |  | -78 | 97.8333 | 31585.8333 | -0.0046593 | 0.00095507 | -0.46593 | 0.0095507 | H | neg | neg
409 | PAMF Pre-COVID | Negative | 581 |  | 95.6667 | 89 | 27141.3333 | 0.00097637 | 0.00078601 | 0.097637 | 0.0078601 | H | neg | neg
410 | PAMF Pre-COVID | Negative | 582 |  | 101 | 163.6667 | 24791.8333 | 0.001284 | 0.0038722 | 0.1284 | 0.038722 | H | neg | neg
411 | PAMF Pre-COVID | Negative | 583 |  | 1 | 67.3333 | 27704.3333 | -0.0024605 | -1.2032e-05 | -0.24605 | -0.00012032 | H | neg | neg
412 | PAMF Pre-COVID | Negative | 584 |  | -0.66667 | 101.3333 | 23645 | -0.0029534 | 0.0014238 | -0.29534 | 0.014238 | H | neg | neg
413 | PAMF Pre-COVID | Negative | 585 |  | 36.3333 | 188.8333 | 25360.5 | -0.0012947 | 0.0047778 | -0.12947 | 0.047778 | H | neg | neg
414 | PAMF Pre-COVID | Negative | 586 |  | -3.6667 | 95.3333 | 25884 | -0.0028138 | 0.0010689 | -0.28138 | 0.010689 | H | neg | neg
415 | PAMF Pre-COVID | Negative | 587 |  | 131 | 91.3333 | 24761.1667 | 0.0024972 | 0.0009558 | 0.24972 | 0.009558 | H | neg | neg
416 | PAMF Pre-COVID | Negative | 588 |  | 104 | 105.6667 | 24160.5 | 0.0014417 | 0.0015728 | 0.14417 | 0.015728 | H | neg | neg
417 | PAMF Pre-COVID | Negative | 589 |  | 97.8333 | 86 | 22928.8333 | 0.0012502 | 0.00079958 | 0.12502 | 0.0079958 | H | neg | neg
418 | PAMF Pre-COVID | Negative | 590 |  | 5.3333 | 101.5 | 21419.5 | -0.0029802 | 0.0015796 | -0.29802 | 0.015796 | H | neg | neg
419 | FL PCR+ | Positive | 201 | NA | 44.6667 | 92 | 14492.1667 | 0.0031051 | 0.0014031 | 0.31051 | 0.014031 | P | neg | neg
420 | FL PCR+ | Positive | 202 | NA | 4.1667 | 91.6667 | 14946.1667 | 0.00030108 | 0.0013381 | 0.030108 | 0.013381 | P | neg | neg
421 | FL PCR+ | Positive | 203 | NA | 676.8333 | 860.5 | 14990.3333 | 0.045174 | 0.052623 | 4.5174 | 0.52623 | P | pos | neg
422 | FL PCR+ | Positive | 204 | NA | 606 | 196.3333 | 13142.5 | 0.041798 | 0.0091053 | 4.1798 | 0.091053 | P | pos | neg
423 | FL PCR+ | Positive | 205 | NA | 3788 | 37818.8333 | 15959.8333 | 0.2338 | 2.3648 | 23.3795 | 23.6482 | P | pos | pos
424 | FL PCR+ | Positive | 206 | NA | 255 | 134.1667 | 14810.1667 | 0.013392 | 0.0038825 | 1.3392 | 0.038825 | P | neg | neg
425 | FL PCR+ | Positive | 207 | NA | 1963 | 3572 | 14327.5 | 0.1367 | 0.24389 | 13.6695 | 2.4389 | P | pos | pos
426 | FL PCR+ | Positive | 208 | NA | 26.3333 | 121.3333 | 14836.1667 | 0.0014716 | 0.0029433 | 0.14716 | 0.029433 | P | neg | neg
427 | FL PCR+ | Positive | 209 | NA | 275.3333 | 1017.3333 | 15087.3333 | 0.017951 | 0.062282 | 1.7951 | 0.62282 | P | pos | neg
428 | FL PCR+ | Positive | 210 | NA | 242.1667 | 70.1667 | 15843 | 0.01518 | -0.00019988 | 1.518 | -0.0019988 | P | pos | neg
429 | FL PCR+ | Positive | 211 | NA | 2.6667 | 99.3333 | 14317 | 6.9847e-05 | 0.001816 | 0.0069847 | 0.01816 | P | neg | neg
430 | FL PCR+ | Positive | 212 | NA | -96.6667 | 105.6667 | 15872.3333 | -0.0061953 | 0.0020371 | -0.61953 | 0.020371 | P | neg | neg
431 | FL PCR+ | Positive | 213 | NA | 700 | 371.1667 | 13586 | 0.051401 | 0.021922 | 5.1401 | 0.21922 | P | pos | neg
432 | FL PCR+ | Positive | 214 | NA | 0 | 0 | 18066.3333 | -9.2253e-05 | -0.0040591 | -0.0092253 | -0.040591 | P | neg | neg
433 | FL PCR+ | Positive | 215 | NA | 103.8333 | 186.6667 | 14863.5 | 0.0068737 | 0.0076249 | 0.68737 | 0.076249 | P | neg | neg
434 | FL PCR+ | Positive | 216 | NA | 87 | 99.3333 | 17120.1667 | 0.0049844 | 0.0015187 | 0.49844 | 0.015187 | P | neg | neg
435 | FL PCR+ | Positive | 217 | NA | 1240.3333 | 35653.3333 | 13015 | 0.095172 | 2.7338 | 9.5172 | 27.3377 | P | pos | pos
436 | FL PCR+ | Positive | 218 | NA | 344.1667 | 109.6667 | 15121 | 0.022651 | 0.0024028 | 2.2651 | 0.024028 | P | pos | neg
437 | FL PCR+ | Positive | 219 | NA | 7.3333 | 98.1667 | 14946.5 | 0.00037913 | 0.0016615 | 0.037913 | 0.016615 | P | neg | neg
438 | FL PCR+ | Positive | 220 | NA | 272.3333 | 149 | 15024 | 0.018016 | 0.0050364 | 1.8016 | 0.050364 | P | pos | neg
439 | FL PCR+ | Positive | 221 | NA | 1870 | 14103.1667 | 14826.3333 | 0.12601 | 0.94628 | 12.6015 | 9.4628 | P | pos | pos
440 | FL PCR+ | Positive | 222 | NA | -37.6667 | 627.6667 | 15310 | -0.0025691 | 0.036207 | -0.25691 | 0.36207 | P | neg | neg
441 | FL PCR+ | Positive | 223 | NA | -16.3333 | 399 | 15271.6667 | -0.0011787 | 0.021325 | -0.11787 | 0.21325 | P | neg | neg
442 | FL PCR+ | Positive | 441 | NA | 82 | 95 | 14994 | -0.0013783 | 0.0007114 | -0.13783 | 0.007114 | P | neg | neg
443 | FL PCR+ | Positive | 442 | NA | 21274.6667 | 65294.6667 | 16989.3333 | 1.2462 | 3.8383 | 124.6194 | 38.3831 | P | pos | pos
444 | FL PCR+ | Positive | 443 | NA | 8539.5 | 754.3333 | 14235.3333 | 0.59267 | 0.047066 | 59.2668 | 0.47066 | P | pos | neg
445 | FL PCR+ | Positive | 444 | NA | 27047 | 65240.6667 | 16036 | 1.6802 | 4.0631 | 168.024 | 40.6313 | P | pos | pos
446 | FL PCR+ | Positive | 445 | NA | 91.3333 | 36.5 | 16647.5 | -0.00068078 | -0.0028733 | -0.068078 | -0.028733 | P | neg | neg
447 | FL PCR+ | Positive | 446 | NA | 98.1667 | 107.3333 | 14262.1667 | -0.00031552 | 0.0016127 | -0.031552 | 0.016127 | P | neg | neg
448 | FL PCR+ | Positive | 447 | NA | 628.3333 | 8017.1667 | 15186.5 | 0.034757 | 0.52295 | 3.4757 | 5.2295 | P | pos | pos
449 | FL PCR+ | Positive | 448 | NA | 3418.6667 | 19725.6667 | 13545.8333 | 0.24496 | 1.4507 | 24.4958 | 14.5066 | P | pos | pos
450 | FL PCR+ | Positive | 449 | NA | 109.3333 | 165.6667 | 15229.1667 | 0.00058003 | 0.0059316 | 0.058003 | 0.059316 | P | neg | neg
451 | FL PCR+ | Positive | 450 | NA | 206.6667 | 101 | 13979.3333 | 0.0075945 | 0.001836 | 0.75945 | 0.01836 | P | neg | neg
452 | FL PCR+ | Positive | 451 | NA | 492.3333 | 6834.1667 | 14333.6667 | 0.027337 | 0.47154 | 2.7337 | 4.7154 | P | pos | pos
453 | FL PCR+ | Positive | 452 | NA | 10591.6667 | 65240.6667 | 13957.8333 | 0.75163 | 4.6687 | 75.1633 | 46.6873 | P | pos | pos
454 | LL PCR+ | Positive | 13028JL | NA | 9355.1667 | 65237 | 14838.8333 | 0.63047 | 4.3915 | 63.0474 | 43.9154 | P | pos | pos
455 | LL PCR+ | Positive | 14902DO | NA | 951.3333 | 19335 | 12521.1667 | 0.076005 | 1.5385 | 7.6005 | 15.3846 | P | pos | pos
456 | LL PCR+ | Positive | 18254AV | NA | 2665.1667 | 65240.6667 | 15528.8333 | 0.17165 | 4.1966 | 17.1648 | 41.9664 | P | pos | pos
457 | LL PCR+ | Positive | 20455WF | NA | 3927 | 65270.3333 | 14740.5 | 0.26643 | 4.4231 | 26.6431 | 44.231 | P | pos | pos
458 | LL PCR+ | Positive | 24897JN | NA | 5445.6667 | 65242 | 14961.6667 | 0.364 | 4.3558 | 36.3997 | 43.5582 | P | pos | pos
459 | LL PCR+ | Positive | 256531IM | NA | 2763.5 | 9806.3333 | 13083.5 | 0.21125 | 0.74404 | 21.1246 | 7.4404 | P | pos | pos
460 | LL PCR+ | Positive | 26437RM | NA | 2812.6667 | 20346.8333 | 14417.8333 | 0.19511 | 1.4063 | 19.5106 | 14.0626 | P | pos | pos
461 | LL PCR+ | Positive | 65363RT | NA | 1783.6667 | 5974.6667 | 13111.5 | 0.13606 | 0.45022 | 13.6064 | 4.5022 | P | pos | pos
462 | LL PCR+ | Positive | 91615MB | NA | 1590.3333 | 46963 | 15350.8333 | 0.10362 | 3.0546 | 10.3621 | 30.5464 | P | pos | pos
463 | LL PCR+ | Positive | 30622DK | NA | 6765.1667 | 65244 | 15302.5 | 0.43839 | 4.2586 | 43.8392 | 42.5861 | P | pos | pos
464 | LL PCR+ | Positive | 32395LP | NA | 3448.6667 | 54595.3333 | 14849 | 0.22843 | 3.6715 | 22.8433 | 36.7154 | P | pos | pos
465 | LL PCR+ | Positive | 32540EJ | NA | 4413 | 65240.3333 | 16595.6667 | 0.2625 | 3.9265 | 26.2498 | 39.2655 | P | pos | pos
466 | LL PCR+ | Positive | 35264HC | NA | 5819 | 61882.3333 | 14725.3333 | 0.39132 | 4.1972 | 39.1321 | 41.9723 | P | pos | pos
467 | LL PCR+ | Positive | 40590DC | NA | 43034 | 65322 | 17451 | 2.4627 | 3.7388 | 246.2743 | 37.3877 | P | pos | pos
468 | LL PCR+ | Positive | 44408AO | NA | 2919.3333 | 2153.6667 | 13840.5 | 0.20683 | 0.15007 | 20.6833 | 1.5007 | P | pos | neg
469 | LL PCR+ | Positive | 57728ER | NA | 967.6667 | 292 | 15563 | 0.058536 | 0.013836 | 5.8536 | 0.13836 | P | pos | neg
470 | LL PCR+ | Positive | 65698LM | NA | 8568.8333 | 65245.6667 | 12826.6667 | 0.66363 | 5.0807 | 66.363 | 50.8074 | P | pos | pos
471 | LL PCR+ | Positive | 34005MH | NA | 5995.8333 | 5504.1667 | 15086.3333 | 0.39368 | 0.35976 | 39.3679 | 3.5976 | P | pos | pos
472 | LL PCR+ | Positive | 45612JS | NA | 11418.3333 | 65302.6667 | 15878.8333 | 0.71881 | 4.1077 | 71.8808 | 41.0767 | P | pos | pos
473 | LL PCR+ | Positive | 90003MS | NA | 459.8333 | 469.8333 | 17034.5 | 0.02673 | 0.023022 | 2.673 | 0.23022 | P | pos | neg
474 | LL PCR+ | Positive | 54225BW | NA | 4022.8333 | 12200 | 15983.1667 | 0.25141 | 0.75844 | 25.141 | 7.5844 | P | pos | pos
475 | LL PCR+ | Positive | 21531NH | NA | 2496.5 | 10171.5 | 17209.8333 | 0.1448 | 0.58652 | 14.4801 | 5.8652 | P | pos | pos
476 | LL PCR+ | Positive | 41279FR | NA | 21261.8333 | 65030 | 18239.5 | 1.1655 | 3.5611 | 116.5456 | 35.6108 | P | pos | pos
477 | LL PCR+ | Positive | 89439MC | NA | 1725.8333 | 2848.3333 | 15939.8333 | 0.10799 | 0.17382 | 10.7989 | 1.7382 | P | pos | pos
478 | LL PCR+ | Positive | 68775KL | NA | 1979.1667 | 998.1667 | 14721.5 | 0.13413 | 0.062528 | 13.4135 | 0.62528 | P | pos | neg
479 | LL PCR+ | Positive | 83211PH | NA | 4309.8333 | 65144.6667 | 14214.1667 | 0.30289 | 4.5776 | 30.289 | 45.7762 | P | pos | pos
480 | LL PCR+ | Positive | 93272RW | NA | 101 | 73.3333 | 14457.8333 | 0.0066746 | -0.00029972 | 0.66746 | -0.0029972 | P | neg | neg
551 | In-House Healthy | Negative | 53 |  | 91.3333 | 102.1667 | 13864.6667 | 0.0019234 | 0.0019113 | 0.19234 | 0.019113 | H | neg | neg
552 | In-House Healthy | Negative | 54 |  | 129.3333 | 17.3333 | 12858.8333 | 0.005029 | -0.0045364 | 0.5029 | -0.045364 | H | neg | neg
553 | In-House Healthy | Negative | 55 |  | 101.6667 | 114 | 13935.3333 | 0.004497 | 0.0025594 | 0.4497 | 0.025594 | H | neg | neg
554 | In-House Healthy | Negative | 56 |  | 95.3333 | 106 | 14598.1667 | 0.0038589 | 0.0018952 | 0.38589 | 0.018952 | H | neg | neg
555 | In-House Healthy | Negative | 57 |  | 154.5 | 76.6667 | 13558 | 0.008519 | -0.00012293 | 0.8519 | -0.0012293 | H | neg | neg
556 | In-House Healthy | Negative | 58 |  | 74.3333 | 69.6667 | 14368 | 0.0024592 | -0.00060319 | 0.24592 | -0.0060319 | H | neg | neg
557 | In-House Healthy | Negative | 59 |  | 97.1667 | 112.6667 | 12990.3333 | 0.0044777 | 0.002643 | 0.44777 | 0.02643 | H | neg | neg
558 | In-House Healthy | Negative | 60 |  | 103.8333 | 110 | 13196.8333 | 0.0044645 | 0.0010987 | 0.44645 | 0.010987 | H | neg | neg
559 | In-House Healthy | Negative | 61 |  | 92 | 75.6667 | 13196.8333 | 0.0044645 | 0.0010987 | 0.44645 | 0.010987 | H | neg | neg
560 | In-House Healthy | Negative | 62 |  | 96.3333 | 73 | 12993.1667 | 0.0044126 | -0.00041047 | 0.44126 | -0.0041047 | H | neg | neg
561 | In-House Healthy | Negative | 63 |  | 93.6667 | 75 | 11573.6667 | 0.0047234 | -0.00028801 | 0.47234 | -0.0028801 | H | neg | neg
562 | In-House Healthy | Negative | 64 |  | 39.6667 | 103.6667 | 12755.5 | 5.2265e-05 | 0.0019861 | 0.0052265 | 0.019861 | H | neg | neg
563 | In-House Healthy | Negative | 65 |  | -227.3333 | 90.8333 | 12019 | -0.022159 | 0.00104 | -2.2159 | 0.0104 | H | neg | neg
564 | In-House Healthy | Negative | 468 |  | 100 | 105.6667 | 16925.3333 | 0.00084686 | 0.0016346 | 0.084686 | 0.016346 | H | neg | neg
565 | In-House Healthy | Negative | 469 |  | 41.1667 | 112.6667 | 16165.6667 | -0.0027527 | 0.0021445 | -0.27527 | 0.021445 | H | neg | neg
566 | In-House Healthy | Negative | 470 |  | 94.3333 | 105.6667 | 16775.8333 | 0.00051662 | 0.0016492 | 0.051662 | 0.016492 | H | neg | neg
567 | In-House Healthy | Negative | 472 |  | 188.3333 | 133.6667 | 17102.1667 | 0.0060031 | 0.0032549 | 0.60031 | 0.032549 | H | neg | neg
568 | In-House Healthy | Negative | 474 |  | 127 | 125.6667 | 18291.1667 | 0.0022597 | 0.002606 | 0.22597 | 0.02606 | H | neg | neg
569 | In-House Healthy | Negative | 475 |  | 95 | 80.6667 | 17888 | 0.00052177 | 0.00014908 | 0.052177 | 0.0014908 | H | neg | neg
570 | In-House Healthy | Negative | 476 |  | 101 | 119 | 17466.1667 | 0.00087789 | 0.0023474 | 0.087789 | 0.023474 | H | neg | neg
571 | In-House Healthy | Negative | 477 |  | 202 | 8.6667 | 15490 | 0.0075102 | -0.004476 | 0.75102 | -0.04476 | H | neg | neg
572 | In-House Healthy | Negative | 479 |  | 99.6667 | 73.6667 | 17196.8333 | 0.0008141 | -0.00025198 | 0.08141 | -0.0025198 | H | neg | neg
573 | In-House Healthy | Negative | 480 |  | 130.6667 | 72.6667 | 16850.8333 | 0.0026705 | -0.0003165 | 0.26705 | -0.003165 | H | neg | neg
574 | In-House Healthy | Negative | 481 |  | 98 | 99 | 15120.5 | 0.00081567 | 0.0013888 | 0.081567 | 0.013888 | H | neg | neg
575 | In-House Healthy | Negative | 485 |  | 154.8333 | 73 | 14973.5 | 0.0046193 | -0.00033392 | 0.46193 | -0.0033392 | H | neg | neg
576 | In-House Healthy | Negative | 487 |  | 68.3333 | 204.1667 | 13101.5 | 0.0022389 | 0.0096045 | 0.22389 | 0.096045 | H | neg | neg
577 | In-House Healthy | Negative | 488 |  | 102.6667 | 109.6667 | 12381.6667 | 0.005142 | 0.0025306 | 0.5142 | 0.025306 | H | neg | neg
578 | In-House Healthy | Negative | 492 |  | -3.3333 | 99.3333 | 12936 | -0.0032725 | 0.0016234 | -0.32725 | 0.016234 | H | neg | neg
579 | In-House Healthy | Negative | 493 |  | 92.6667 | 108.1667 | 16350.6667 | 0.0017125 | 0.0019877 | 0.17125 | 0.019877 | H | neg | neg
580 | In-House Healthy | Negative | 495 |  | 193.3333 | 103.8333 | 15259.5 | 0.0084319 | 0.0018458 | 0.84319 | 0.018458 | H | neg | neg
581 | In-House Healthy | Negative | 498 |  | 8.6667 | 103.3333 | 14628.3333 | -0.0038282 | 0.0018913 | -0.38282 | 0.018913 | H | neg | neg
582 | In-House Healthy | Negative | 499 |  | 102.3333 | 103.6667 | 15909.3333 | 0.0023676 | 0.00176 | 0.23676 | 0.0176 | H | neg | neg
583 | In-House Healthy | Negative | 500 |  | 100.1667 | 197.3333 | 14840.5 | 0.0023921 | 0.0081983 | 0.23921 | 0.081983 | H | neg | neg
584 | In-House Healthy | Negative | 501 |  | 115 | 103.5 | 14346.3333 | 0.0035084 | 0.0019401 | 0.35084 | 0.019401 | H | neg | neg
585 | In-House Healthy | Negative | 505 |  | 66.6667 | 73.5 | 13757.6667 | 0.00014537 | -0.00015749 | 0.014537 | -0.0015749 | H | neg | neg
586 | In-House Healthy | Negative | 507 |  | 180.6667 | 57.6667 | 14727.6667 | 0.0078763 | -0.0012222 | 0.78763 | -0.012222 | H | neg | neg
587 | In-House Healthy | Negative | 508 |  | 96.3333 | 96.3333 | 13926.6667 | 0.0022738 | 0.001484 | 0.22738 | 0.01484 | H | neg | neg
588 | In-House Healthy | Negative | 510 |  | 105 | 106 | 12921.5 | 0.0031214 | 0.0023475 | 0.31214 | 0.023475 | H | neg | neg
589 | In-House Healthy | Negative | 551 |  | 65.6667 | 11 | 12752.5 | 7.8416e-05 | -0.0050709 | 0.0078416 | -0.050709 | H | neg | neg
590 | In-House Healthy | Negative | 552 |  | -85 | 13 | 12302.3333 | -0.012166 | -0.0050939 | -1.2166 | -0.050939 | H | neg | neg

## ALL DATA
Sheet_Order | Source/Sample Type | Study Subcategory | Expected Result | Patient ID | CR: Disease Type | PCR+: Days Since Symptom Onset | IgM MFI | IgG MFI | Intrawell Standard MFI | IgM Intrawell Ratio | IgG Intrawell Ratio | IgM Intrawell, adjusted | IgG Intrawell, adjusted | Status | S1 Intrawell_IgM (<=1.38) | S1 Intrawell_IgG (<=1.62)
1 | PAMF PCR+ | PCR+ | Positive | 1 |  | NA | 62 | 49 | 19804.6667 | -0.001296 | -0.001397 | -0.1296 | -0.01397 | P | neg | neg
2 | PAMF PCR+ | PCR+ | Positive | 2 |  | NA | 158.1667 | 111 | 17690.6667 | 0.0039852 | 0.0019408 | 0.39852 | 0.019408 | P | neg | neg
3 | PAMF PCR+ | PCR+ | Positive | 3 |  | NA | 30 | 18.3333 | 16950.5 | -0.0034021 | -0.0034414 | -0.34021 | -0.034414 | P | neg | neg
4 | PAMF PCR+ | PCR+ | Positive | 4 |  | 4 | 93.3333 | 164 | 17655.6667 | 0.00032095 | 0.0049465 | 0.032095 | 0.049465 | P | neg | neg
5 | PAMF PCR+ | PCR+ | Positive | 5 |  | 1 | 6.3333 | 100.8333 | 16008.8333 | -0.0050805 | 0.0015096 | -0.50805 | 0.015096 | P | neg | neg
6 | PAMF PCR+ | PCR+ | Positive | 6 |  | 5 | 32.6667 | 101.6667 | 16923.3333 | -0.00325 | 0.0014773 | -0.325 | 0.014773 | P | neg | neg
7 | PAMF PCR+ | PCR+ | Positive | 7 |  | 5 | 5.6667 | 187 | 15618.5 | -0.0052502 | 0.0070643 | -0.52502 | 0.070643 | P | neg | neg
8 | PAMF PCR+ | PCR+ | Positive | 8 |  | 0 | -2.8333 | 149 | 15909.8333 | -0.0056883 | 0.0045465 | -0.56883 | 0.045465 | P | neg | neg
9 | PAMF PCR+ | PCR+ | Positive | 9 |  | 23 | 8095.6667 | 65270.6667 | 15799.3333 | 0.50686 | 4.1264 | 50.6857 | 41.2638 | P | pos | pos
10 | PAMF PCR+ | PCR+ | Positive | 10 |  | 19 | 5360.1667 | 65264.6667 | 16662.5 | 0.31643 | 3.9123 | 31.6429 | 39.1226 | P | pos | pos
11 | PAMF PCR+ | PCR+ | Positive | 11 |  | 23 | 920 | 20911.3333 | 16020.5 | 0.051954 | 1.3005 | 5.1954 | 13.005 | P | pos | pos
12 | PAMF PCR+ | PCR+ | Positive | 12 |  | 2 | 855.8333 | 1093.8333 | 16532.1667 | 0.046465 | 0.061527 | 4.6465 | 0.61527 | P | pos | neg
13 | PAMF PCR+ | PCR+ | Positive | 13 |  | 18 | 3149.8333 | 40129.3333 | 17521 | 0.17477 | 2.286 | 17.4771 | 22.8598 | P | pos | pos
14 | PAMF PCR+ | PCR+ | Positive | 14 |  | 12 | 499 | 2786.1667 | 18879.1667 | 0.027526 | 0.1435 | 2.7526 | 1.435 | P | pos | neg
15 | PAMF PCR+ | PCR+ | Positive | 15 |  | 7 | 1316.6667 | 2047.1667 | 18282.1667 | 0.070114 | 0.1074 | 7.0114 | 1.074 | P | pos | neg
16 | PAMF PCR+ | PCR+ | Positive | 16 |  | 19 | 8222 | 65149.3333 | 18689.1667 | 0.43807 | 3.4815 | 43.807 | 34.8146 | P | pos | pos
17 | PAMF PCR+ | PCR+ | Positive | 17 |  | 3 | 1543.5 | 3067.6667 | 19896.8333 | 0.078614 | 0.15031 | 7.8614 | 1.5031 | P | pos | neg
18 | PAMF PCR+ | PCR+ | Positive | 18 |  | 38 | 16406.3333 | 65236.8333 | 17996 | 0.90973 | 3.6204 | 90.973 | 36.2042 | P | pos | pos
19 | PAMF PCR+ | PCR+ | Positive | 19 |  | NA | 2475.3333 | 1035.6667 | 17490.1667 | 0.13954 | 0.054431 | 13.9536 | 0.54431 | P | pos | neg
20 | PAMF PCR+ | PCR+ | Positive | 20 |  | 1 | -55.1667 | 192.5 | 19524.1667 | -0.001767 | 0.0059157 | -0.1767 | 0.059157 | P | neg | neg
21 | PAMF PCR+ | PCR+ | Positive | 21 |  | 4 | 3.3333 | 311.3333 | 20449.5 | 0.0011736 | 0.011459 | 0.11736 | 0.11459 | P | neg | neg
22 | PAMF PCR+ | PCR+ | Positive | 22 |  | 23 | 1083.8333 | 19880.3333 | 16434.8333 | 0.063828 | 1.2046 | 6.3828 | 12.0456 | P | pos | pos
23 | PAMF PCR+ | PCR+ | Positive | 64 |  | 8 | -2 | 167.3333 | 16035 | -0.0044694 | 0.0056335 | -0.44694 | 0.056335 | P | neg | neg
24 | PAMF PCR+ | PCR+ | Positive | 65 |  | 6 | 437 | 28797 | 18331.5 | 0.020038 | 1.5667 | 2.0038 | 15.667 | P | pos | pos
25 | PAMF PCR+ | PCR+ | Positive | 70 |  | 9 | 104 | 124.3333 | 18002.6667 | 0.0055547 | 0.0027218 | 0.55547 | 0.027218 | P | neg | neg
26 | PAMF PCR+ | PCR+ | Positive | 73 |  | 12 | 459.3333 | 2003.3333 | 14068 | 0.027699 | 0.13693 | 2.7699 | 1.3693 | P | pos | neg
27 | PAMF PCR+ | PCR+ | Positive | 74 |  | 7 | 629.3333 | 1782 | 16969.6667 | 0.03298 | 0.10047 | 3.298 | 1.0047 | P | pos | neg
28 | PAMF PCR+ | PCR+ | Positive | 78 |  | 10 | 197.3333 | 363 | 15652.5 | 0.012352 | 0.018378 | 1.2352 | 0.18378 | P | neg | neg
29 | PAMF PCR+ | PCR+ | Positive | 100 |  | 7 | 102.3333 | 101 | 18121.6667 | 0.0054263 | 0.0014164 | 0.54263 | 0.014164 | P | neg | neg
30 | PAMF PCR+ | PCR+ | Positive | 101 |  | 13 | 1033.5 | 19406.8333 | 18710 | 0.051514 | 1.0331 | 5.1514 | 10.3313 | P | pos | pos
31 | PAMF PCR+ | PCR+ | Positive | 102 |  | 8 | 382.5 | 1625.1667 | 15662 | 0.019974 | 0.098849 | 1.9974 | 0.98849 | P | pos | neg
32 | PAMF PCR+ | PCR+ | Positive | 114 |  | 14 | 199.5 | 434.3333 | 17361 | 0.011261 | 0.020679 | 1.1261 | 0.20679 | P | neg | neg
33 | PAMF PCR+ | PCR+ | Positive | 115 |  | 13 | -0.33333 | 27.6667 | 18687.3333 | -0.00023189 | -0.0025507 | -0.023189 | -0.025507 | P | neg | neg
34 | CDPH PCR- | PCR- | Negative | 2.1 |  |  | 2.3333 | 99 | 15217.6667 | -0.0034171 | 0.0015771 | -0.34171 | 0.015771 | H | neg | neg
35 | CDPH PCR- | PCR- | Negative | 2.2 |  |  | -9 | 194.3333 | 18417.5 | -0.0040994 | 0.0063889 | -0.40994 | 0.063889 | H | neg | neg
36 | CDPH PCR- | PCR- | Negative | 2.3 |  |  | -4.6667 | 186.3333 | 17667.3333 | -0.0040281 | 0.0062073 | -0.40281 | 0.062073 | H | neg | neg
37 | CDPH PCR- | PCR- | Negative | 2.4 |  |  | -1.3333 | 106.6667 | 16914.6667 | -0.0040103 | 0.0017736 | -0.40103 | 0.017736 | H | neg | neg
38 | CDPH PCR+ | PCR+ | Positive | 2.18 |  | 12 | 64938.6667 | 64031.6667 | 15411.3333 | 4.2135 | 4.1497 | 421.3544 | 41.4967 | P | pos | pos
39 | CDPH PCR+ | PCR+ | Positive | 2.19 |  | 23 | 11951 | 65201.1667 | 17547.8333 | 0.68092 | 3.7111 | 68.092 | 37.1108 | P | pos | pos
40 | CDPH PCR+ | PCR+ | Positive | 2.2 |  | 22 | 12135.1667 | 64919.6667 | 17673.6667 | 0.68649 | 3.6687 | 68.6492 | 36.6874 | P | pos | pos
41 | CDPH PCR+ | PCR+ | Positive | 2.21 |  | 17 | 23576 | 65179 | 15056 | 1.5657 | 4.3238 | 156.5732 | 43.2381 | P | pos | pos
42 | CDPH PCR+ | PCR+ | Positive | 2.22 |  | 17 | 21542.1667 | 65240.6667 | 15650.8333 | 1.3763 | 4.1634 | 137.6274 | 41.6342 | P | pos | pos
43 | CDPH PCR+ | PCR+ | Positive | 2.24 |  | 20 | 14376 | 65230.6667 | 17032.3333 | 0.84085 | 3.8254 | 84.0852 | 38.2541 | P | pos | pos
44 | CDPH PCR+ | PCR+ | Positive | 2.25 |  | 21 | 11086.1667 | 65161.6667 | 19299.6667 | 0.57161 | 3.3724 | 57.1607 | 33.7242 | P | pos | pos
45 | CDPH PCR+ | PCR+ | Positive | 2.26 |  | 19 | 10808.3333 | 65236.3333 | 18010.5 | 0.5971 | 3.618 | 59.7096 | 36.1796 | P | pos | pos
46 | CDPH PCR+ | PCR+ | Positive | 2.27 |  | 17 | 4766 | 64833 | 17714.5 | 0.26598 | 3.6556 | 26.5978 | 36.5565 | P | pos | pos
47 | CDPH PCR+ | PCR+ | Positive | 2.28 |  | 12 | 2654.1667 | 22406.3333 | 16335.3333 | 0.15915 | 1.3671 | 15.9154 | 13.6706 | P | pos | pos
48 | CDPH PCR+ | PCR+ | Positive | 2.29 |  | 12 | 4193.3333 | 59179.6667 | 16325.3333 | 0.25353 | 3.6204 | 25.3532 | 36.2043 | P | pos | pos
49 | CDPH PCR+ | PCR+ | Positive | 2.3 |  | 20 | 10179.1667 | 65149.6667 | 15838.6667 | 0.63925 | 4.1086 | 63.9248 | 41.0859 | P | pos | pos
50 | CDPH PCR+ | PCR+ | Positive | 2.31 |  | 18 | 5053.8333 | 64734.3333 | 14103.3333 | 0.35449 | 4.5847 | 35.4491 | 45.8468 | P | pos | pos
51 | CDPH PCR+ | PCR+ | Positive | 2.32 |  | 11 | 1112.8333 | 5603.6667 | 15196.6667 | 0.069653 | 0.36381 | 6.9653 | 3.6381 | P | pos | pos
52 | CDPH PCR+ | PCR+ | Positive | 2.35 |  | 22 | 8129.6667 | 65130.3333 | 15090.6667 | 0.53512 | 4.311 | 53.5121 | 43.1096 | P | pos | pos
53 | CDPH PCR+ | PCR+ | Positive | 2.36 |  | 18 | 22001.6667 | 65147.3333 | 13935.5 | 1.5749 | 4.6695 | 157.4923 | 46.6954 | P | pos | pos
54 | CDPH PCR+ | PCR+ | Positive | 2.38 |  | 25 | 6850.5 | 65152.3333 | 15247.3333 | 0.44573 | 4.2681 | 44.5728 | 42.6811 | P | pos | pos
55 | CDPH PCR+ | PCR+ | Positive | 2.4 |  | 13 | 490.8333 | 2428.6667 | 14073.1667 | 0.031016 | 0.16724 | 3.1016 | 1.6724 | P | pos | pos
56 | CDPH PCR+ | PCR+ | Positive | 2 |  | 10 | 4519.1667 | 12411.8333 | 18287.8333 | 0.24521 | 0.67412 | 24.5209 | 6.7412 | P | pos | pos
57 | CDPH PCR+ | PCR+ | Positive | 2.1 |  | 10 | 950.3333 | 444.8333 | 15182.6667 | 0.062637 | 0.024425 | 6.2637 | 0.24425 | P | pos | neg
58 | CDPH PCR+ | PCR+ | Positive | 2.11 |  | 32 | 7006.8333 | 64604.6667 | 15705.3333 | 0.44619 | 4.1088 | 44.6186 | 41.0884 | P | pos | pos
59 | CDPH PCR+ | PCR+ | Positive | 2.12 |  | 26 | 10098.3333 | 63696.1667 | 18070.1667 | 0.55888 | 3.5208 | 55.8877 | 35.2084 | P | pos | pos
60 | CDPH PCR+ | PCR+ | Positive | 2.13 |  | 20 | 41016.3333 | 65236.6667 | 17413.1667 | 2.3555 | 3.7421 | 235.5516 | 37.4215 | P | pos | pos
61 | CDPH PCR+ | PCR+ | Positive | 2.14 |  | 27 | 3653.3333 | 65144 | 16678.3333 | 0.21909 | 3.9015 | 21.9087 | 39.0147 | P | pos | pos
62 | CDPH PCR+ | PCR+ | Positive | 2.15 |  | 36 | 5750.8333 | 65244 | 16829.1667 | 0.34176 | 3.8724 | 34.1758 | 38.7244 | P | pos | pos
63 | CDPH PCR+ | PCR+ | Positive | 2.16 |  | 27 | 12840.8333 | 65139.3333 | 16108.3333 | 0.7972 | 4.0392 | 79.7196 | 40.3923 | P | pos | pos
64 | CDPH PCR+ | PCR+ | Positive | 2.17 |  | 27 | 13382.1667 | 65259.6667 | 17448.8333 | 0.76698 | 3.7358 | 76.6976 | 37.3582 | P | pos | pos
65 | CDPH PCR+ | PCR+ | Positive | 2.6 |  | 15 | 12411.3333 | 55430 | 16409 | 0.75641 | 3.3735 | 75.6414 | 33.7351 | P | pos | pos
66 | CDPH PCR+ | PCR+ | Positive | 2.7 |  | 26 | 9626.6667 | 65229.6667 | 16955.3333 | 0.56781 | 3.8428 | 56.7806 | 38.4278 | P | pos | pos
67 | CDPH PCR+ | PCR+ | Positive | 2.8 |  | 16 | 11989.8333 | 65240.3333 | 15814.6667 | 0.75819 | 4.1206 | 75.8189 | 41.2063 | P | pos | pos
68 | CDPH PCR+ | PCR+ | Positive | 2.9 |  | 45 | 4873.6667 | 65237.6667 | 15358.1667 | 0.31738 | 4.2429 | 31.7377 | 42.4293 | P | pos | pos
69 | CDPH PCR+ | PCR+ | Positive | 5 |  | 13 | 3296.6667 | 20004.5 | 18404.3333 | 0.17723 | 1.0824 | 17.7232 | 10.824 | P | pos | pos
70 | CDPH PCR+ | PCR+ | Positive | 7 |  | 15 | 11307.8333 | 65178.6667 | 17590 | 0.64088 | 3.7007 | 64.0875 | 37.0068 | P | pos | pos
71 | CDPH PCR+ | PCR+ | Positive | 9 |  | 15 | 65037 | 65268.3333 | 18618 | 3.4914 | 3.5012 | 349.1361 | 35.0116 | P | pos | pos
72 | CDPH PCR+ | PCR+ | Positive | 14 |  | 14 | 65037.6667 | 65175.6667 | 17580.5 | 3.6974 | 3.7025 | 369.7439 | 37.0251 | P | pos | pos
73 | CDPH PCR+ | PCR+ | Positive | 15 |  | 14 | 2799 | 34781.1667 | 17557.3333 | 0.15744 | 1.9762 | 15.7437 | 19.7624 | P | pos | pos
74 | CDPH PCR+ | PCR+ | Positive | 6 |  | 9 | 201 | 613.6667 | 14680.6667 | 0.013671 | 0.036631 | 1.3671 | 0.36631 | P | neg | neg
75 | CDPH PCR+ | PCR+ | Positive | 8 |  | 2 | 37 | 107 | 13867.3333 | 0.0026449 | 0.0022359 | 0.26449 | 0.022359 | P | neg | neg
76 | CDPH PCR+ | PCR+ | Positive | 16 |  | 8 | 109.1667 | 132 | 12257.8333 | 0.0092166 | 0.0044861 | 0.92166 | 0.044861 | P | neg | neg
77 | CDPH PCR+ | PCR+ | Positive | 2.34 |  | 6 | 357.3333 | 278.1667 | 13470.8333 | 0.025977 | 0.014312 | 2.5977 | 0.14312 | P | pos | neg
78 | CDPH PCR+ | PCR+ | Positive | 2.37 |  | 7 | 5691.6667 | 17528.5 | 15699.1667 | 0.36226 | 1.1111 | 36.2258 | 11.1106 | P | pos | pos
79 | LL PCR- | PCR- | Negative | 06872ER |  |  | 0.16667 | 146.1667 | 17672.6667 | -0.0037534 | 0.0039326 | -0.37534 | 0.039326 | H | neg | neg
80 | LL PCR- | PCR- | Negative | 98671DR |  |  | 3.6667 | 1135.3333 | 16660.6667 | -0.0037714 | 0.063543 | -0.37714 | 0.63543 | H | neg | neg
81 | LL PCR- | PCR- | Negative | 68409DP |  |  | -0.5 | 95.8333 | 16952.6667 | -0.0039522 | 0.0011306 | -0.39522 | 0.011306 | H | neg | neg
82 | LL PCR- | PCR- | Negative | 52104RH |  |  | 3 | 95.6667 | 17116 | -0.00371 | 0.0011101 | -0.371 | 0.011101 | H | neg | neg
83 | LL PCR- | PCR- | Negative | 33492KM |  |  | -3.3333 | 94 | 16197 | -0.0043115 | 0.0010702 | -0.43115 | 0.010702 | H | neg | neg
84 | LL PCR- | PCR- | Negative | 73605JA |  |  | -25.8333 | 207.8333 | 16552 | -0.0055784 | 0.0079245 | -0.55784 | 0.079245 | H | neg | neg
85 | LL PCR- | PCR- | Negative | 27215JD |  |  | 5 | 95 | 16476.6667 | -0.0037326 | 0.0011127 | -0.37326 | 0.011127 | H | neg | neg
86 | LL PCR- | PCR- | Negative | 27829JP |  |  | 41.6667 | 113.8333 | 16551 | -0.0015004 | 0.0022456 | -0.15004 | 0.022456 | H | neg | neg
87 | LL PCR- | PCR- | Negative | 08751EC |  |  | -1 | 97.3333 | 17456.3333 | -0.0038668 | 0.0011839 | -0.38668 | 0.011839 | H | neg | neg
88 | LL PCR- | PCR- | Negative | 05458WP |  |  | 2.3333 | 99.3333 | 16660.5 | -0.0038514 | 0.0013605 | -0.38514 | 0.013605 | H | neg | neg
89 | LL PCR- | PCR- | Negative | 45500BH |  |  | 1.6667 | 449.1667 | 17867 | -0.0036287 | 0.020848 | -0.36287 | 0.20848 | H | neg | neg
90 | LL PCR- | PCR- | Negative | 09742JA |  |  | 6 | 168.6667 | 18620.8333 | 0.0014321 | 0.0049228 | 0.14321 | 0.049228 | H | neg | neg
91 | LL PCR- | PCR- | Negative | 20443PM |  |  | -718.6667 | 133 | 18367.1667 | -0.038003 | 0.0030489 | -3.8003 | 0.030489 | H | neg | neg
92 | LL PCR- | PCR- | Negative | 66311AC |  |  | 2.6667 | 67.8333 | 16714.5 | 0.00019943 | -0.00036894 | 0.019943 | -0.0036894 | H | neg | neg
93 | LL PCR- | PCR- | Negative | 42089SR |  |  | 99.1667 | 136.8333 | 15572.3333 | 0.0064109 | 0.0040349 | 0.64109 | 0.040349 | H | neg | neg
94 | LL PCR- | PCR- | Negative | 02065SJ |  |  | 24 | 67.1667 | 17770 | 0.0011255 | -0.00045958 | 0.11255 | -0.0045958 | H | neg | neg
95 | LL PCR- | PCR- | Negative | 46577ME |  |  | 6 | 103 | 17217.6667 | 0.00011616 | 0.0016069 | 0.011616 | 0.016069 | H | neg | neg
96 | LL PCR- | PCR- | Negative | 07018AA |  |  | 63.5 | 109.6667 | 17985 | 0.0033824 | 0.002984 | 0.33824 | 0.02984 | H | neg | neg
97 | LL PCR- | PCR- | Negative | 56891ER |  |  | -0.66667 | 72.3333 | 17840 | -0.00018685 | 0.00091555 | -0.018685 | 0.0091555 | H | neg | neg
98 | LL PCR- | PCR- | Negative | 42334BS |  |  | -2.5 | 191.3333 | 19103 | -0.00027046 | 0.0070844 | -0.027046 | 0.070844 | H | neg | neg
99 | LL PCR- | PCR- | Negative | 60812KW |  |  | 76.3333 | 106.5 | 18412 | 0.004001 | 0.0027428 | 0.4001 | 0.027428 | H | neg | neg
100 | LL PCR- | PCR- | Negative | 54431LR |  |  | 77.5 | 66.6667 | 15936.8333 | 0.00049152 | -0.00064839 | 0.049152 | -0.0064839 | H | neg | neg
101 | LL PCR- | PCR- | Negative | 30056LC |  |  | 204.3333 | 91.3333 | 15164.3333 | 0.0088805 | 0.0009452 | 0.88805 | 0.009452 | H | neg | neg
102 | LL PCR- | PCR- | Negative | 76896SA |  |  | 3.3333 | 69 | 19046.5 | 3.5002e-05 | 0.00068254 | 0.0035002 | 0.0068254 | H | neg | neg
103 | LL PCR- | PCR- | Negative | 28570MR |  |  | 95.6667 | 108.1667 | 17212.5 | 0.0054031 | 0.0030307 | 0.54031 | 0.030307 | H | neg | neg
104 | LL PCR- | PCR- | Negative | 24087RS |  |  | 35.6667 | 34.3333 | 17824.3333 | 0.0018514 | -0.0012156 | 0.18514 | -0.012156 | H | neg | neg
105 | LL PCR- | PCR- | Negative | 25431LY |  |  | 56 | 7.6667 | 16335.5 | 0.0032649 | -0.0029588 | 0.32649 | -0.029588 | H | neg | neg
106 | LL PCR- | PCR- | Negative | 08286QA |  |  | -5.3333 | 103 | 16198.1667 | -0.00049388 | 0.0029016 | -0.049388 | 0.029016 | H | neg | neg
107 | LL PCR- | PCR- | Negative | 62379AD |  |  | 105.6667 | 105.8333 | 16364.6667 | 0.006294 | 0.0030452 | 0.6294 | 0.030452 | H | neg | neg
108 | PAMF Pre-COVID | Pre-Pandemic | Negative | 2 |  |  | 127.8333 | 102.3333 | 13574.3333 | 0.0068634 | 0.0035361 | 0.68634 | 0.035361 | H | neg | neg
109 | PAMF Pre-COVID | Pre-Pandemic | Negative | 8 |  |  | 2.3333 | 204.6667 | 12637.1667 | -0.0025586 | 0.011896 | -0.25586 | 0.11896 | H | neg | neg
110 | PAMF Pre-COVID | Pre-Pandemic | Negative | 13 |  |  | 103 | 191 | 13037.1667 | 0.0052414 | 0.010483 | 0.52414 | 0.10483 | H | neg | neg
111 | PAMF Pre-COVID | Pre-Pandemic | Negative | 14 |  |  | 24 | 113.1667 | 15773.1667 | -0.00038039 | 0.001511 | -0.038039 | 0.01511 | H | neg | neg
112 | PAMF Pre-COVID | Pre-Pandemic | Negative | 18 |  |  | -0.66667 | 78.3333 | 15728.5 | -0.0019498 | -0.00069937 | -0.19498 | -0.0069937 | H | neg | neg
113 | PAMF Pre-COVID | Pre-Pandemic | Negative | 22 |  |  | -1.8333 | 85 | 17106 | -0.0018609 | -0.00025332 | -0.18609 | -0.0025332 | H | neg | neg
114 | PAMF Pre-COVID | Pre-Pandemic | Negative | 23 |  |  | 2.3333 | 117.3333 | 15504 | -0.0017845 | 0.001806 | -0.17845 | 0.01806 | H | neg | neg
115 | PAMF Pre-COVID | Pre-Pandemic | Negative | 29 |  |  | 34.8333 | 143.8333 | 15466.1667 | 0.00031251 | 0.0035238 | 0.031251 | 0.035238 | H | neg | neg
116 | PAMF Pre-COVID | Pre-Pandemic | Negative | 33 |  |  | 103.8333 | 188 | 14617.1667 | 0.0050511 | 0.0067501 | 0.50511 | 0.067501 | H | neg | neg
117 | PAMF Pre-COVID | Pre-Pandemic | Negative | 81 |  |  | 31.3333 | 105.8333 | 16307.6667 | 8.1761e-05 | 0.0010118 | 0.0081761 | 0.010118 | H | neg | neg
118 | PAMF Pre-COVID | Pre-Pandemic | Negative | 84 |  |  | 37 | 147.5 | 15768 | 0.00044394 | 0.0036889 | 0.044394 | 0.036889 | H | neg | neg
119 | PAMF Pre-COVID | Pre-Pandemic | Negative | 87 |  |  | 5.3333 | 111.3333 | 15810.1667 | -0.0015602 | 0.0013915 | -0.15602 | 0.013915 | H | neg | neg
120 | PAMF Pre-COVID | Pre-Pandemic | Negative | 88 |  |  | 2 | 104.3333 | 15570.6667 | -0.0017983 | 0.00096335 | -0.17983 | 0.0096335 | H | neg | neg
121 | PAMF Pre-COVID | Pre-Pandemic | Negative | 200 |  |  | -104.67 | 35.67 | 21178.33 | -0.00512 | 0.00022 | -0.512 | 0.0022 | H | neg | neg
122 | PAMF Pre-COVID | Pre-Pandemic | Negative | 201 |  |  | -0.83 | 41.17 | 16851.5 | -0.00027 | 0.0006 | -0.027 | 0.006 | H | neg | neg
123 | PAMF Pre-COVID | Pre-Pandemic | Negative | 202 |  |  | -42.17 | 65.33 | 17560.17 | -0.00261 | 0.00196 | -0.261 | 0.0196 | H | neg | neg
124 | PAMF Pre-COVID | Pre-Pandemic | Negative | 203 |  |  | -34 | 101.5 | 16089.83 | -0.00234 | 0.00438 | -0.234 | 0.0438 | H | neg | neg
125 | PAMF Pre-COVID | Pre-Pandemic | Negative | 204 |  |  | -1.17 | 543 | 17505.5 | -0.00028 | 0.02925 | -0.028 | 0.2925 | H | neg | neg
126 | PAMF Pre-COVID | Pre-Pandemic | Negative | 205 |  |  | 5 | 98.33 | 15260.17 | 9e-05 | 0.00441 | 0.009 | 0.0441 | H | neg | neg
127 | PAMF Pre-COVID | Pre-Pandemic | Negative | 206 |  |  | 1.33 | 96.67 | 16554 | -0.00014 | 0.00397 | -0.014 | 0.0397 | H | neg | neg
128 | PAMF Pre-COVID | Pre-Pandemic | Negative | 207 |  |  | -5 | 94 | 16982.67 | -0.00051 | 0.00371 | -0.051 | 0.0371 | H | neg | neg
129 | PAMF Pre-COVID | Pre-Pandemic | Negative | 208 |  |  | 0.83 | 89.83 | 16320.67 | -0.00017 | 0.0036 | -0.017 | 0.036 | H | neg | neg
130 | PAMF Pre-COVID | Pre-Pandemic | Negative | 209 |  |  | -1.5 | 91.33 | 17744.5 | -0.00029 | 0.0034 | -0.029 | 0.034 | H | neg | neg
131 | PAMF Pre-COVID | Pre-Pandemic | Negative | 210 |  |  | -2.33 | 97.33 | 16194.67 | -0.00037 | 0.0041 | -0.037 | 0.041 | H | neg | neg
132 | PAMF Pre-COVID | Pre-Pandemic | Negative | 211 |  |  | -1.33 | 97.33 | 15641.83 | -0.00032 | 0.00424 | -0.032 | 0.0424 | H | neg | neg
133 | PAMF Pre-COVID | Pre-Pandemic | Negative | 212 |  |  | -106.33 | 6 | 22540.33 | -0.006 | -0.00314 | -0.6 | -0.0314 | H | neg | neg
134 | PAMF Pre-COVID | Pre-Pandemic | Negative | 213 |  |  | 5 | 112.17 | 21591.5 | -0.0011 | 0.00164 | -0.11 | 0.0164 | H | neg | neg
135 | PAMF Pre-COVID | Pre-Pandemic | Negative | 214 |  |  | -392 | 29.17 | 20342 | -0.02069 | -0.00234 | -2.069 | -0.0234 | H | neg | neg
136 | PAMF Pre-COVID | Pre-Pandemic | Negative | 215 |  |  | -27.33 | 43 | 20119.5 | -0.00279 | -0.00167 | -0.279 | -0.0167 | H | neg | neg
137 | PAMF Pre-COVID | Pre-Pandemic | Negative | 216 |  |  | 3.33 | 101.67 | 19878.33 | -0.00128 | 0.00126 | -0.128 | 0.0126 | H | neg | neg
138 | PAMF Pre-COVID | Pre-Pandemic | Negative | 217 |  |  | 70.33 | 101.33 | 19927.5 | 0.00208 | 0.00124 | 0.208 | 0.0124 | H | neg | neg
139 | PAMF Pre-COVID | Pre-Pandemic | Negative | 218 |  |  | -4.33 | 9.5 | 20779.5 | -0.0016 | -0.00323 | -0.16 | -0.0323 | H | neg | neg
140 | PAMF Pre-COVID | Pre-Pandemic | Negative | 219 |  |  | -3.33 | 97.67 | 21724.17 | -0.00148 | 0.00097 | -0.148 | 0.0097 | H | neg | neg
141 | PAMF Pre-COVID | Pre-Pandemic | Negative | 220 |  |  | -2.33 | 16.5 | 20032.67 | -0.00156 | -0.003 | -0.156 | -0.03 | H | neg | neg
142 | PAMF Pre-COVID | Pre-Pandemic | Negative | 221 |  |  | 63.33 | 107.67 | 22483 | 0.00153 | 0.00138 | 0.153 | 0.0138 | H | neg | neg
143 | PAMF Pre-COVID | Pre-Pandemic | Negative | 222 |  |  | -0.33 | 146.83 | 18249.17 | -0.0016 | 0.00384 | -0.16 | 0.0384 | H | neg | neg
144 | PAMF Pre-COVID | Pre-Pandemic | Negative | 223 |  |  | 2.33 | 102.5 | 20074.33 | -0.00132 | 0.00129 | -0.132 | 0.0129 | H | neg | neg
145 | PAMF Pre-COVID | Pre-Pandemic | Negative | 224 |  |  | -45.67 | 101 | 21384.67 | -0.00228 | 0.00114 | -0.228 | 0.0114 | H | neg | neg
146 | PAMF Pre-COVID | Pre-Pandemic | Negative | 225 |  |  | -32.67 | 66.67 | 19855.83 | -0.0018 | -0.0005 | -0.18 | -0.005 | H | neg | neg
147 | PAMF Pre-COVID | Pre-Pandemic | Negative | 226 |  |  | -48.5 | 95.67 | 19735 | -0.00261 | 0.00096 | -0.261 | 0.0096 | H | neg | neg
148 | PAMF Pre-COVID | Pre-Pandemic | Negative | 227 |  |  | -9.67 | 100 | 19438.67 | -0.00065 | 0.0012 | -0.065 | 0.012 | H | neg | neg
149 | PAMF Pre-COVID | Pre-Pandemic | Negative | 228 |  |  | -103.83 | 110 | 18198 | -0.00587 | 0.00183 | -0.587 | 0.0183 | H | neg | neg
150 | PAMF Pre-COVID | Pre-Pandemic | Negative | 229 |  |  | 2.33 | 298.33 | 18780.83 | -4e-05 | 0.0118 | -0.004 | 0.118 | H | neg | neg
151 | PAMF Pre-COVID | Pre-Pandemic | Negative | 230 |  |  | 9.67 | 2 | 19837.83 | 0.00034 | -0.00376 | 0.034 | -0.0376 | H | neg | neg
152 | PAMF Pre-COVID | Pre-Pandemic | Negative | 231 |  |  | -0.5 | 91.67 | 18954.17 | -0.00018 | 0.00079 | -0.018 | 0.0079 | H | neg | neg
153 | PAMF Pre-COVID | Pre-Pandemic | Negative | 232 |  |  | -3 | 102 | 17700.33 | -0.00034 | 0.00143 | -0.034 | 0.0143 | H | neg | neg
154 | PAMF Pre-COVID | Pre-Pandemic | Negative | 233 |  |  | 28.33 | 142 | 20515.5 | 0.00123 | 0.00318 | 0.123 | 0.0318 | H | neg | neg
155 | PAMF Pre-COVID | Pre-Pandemic | Negative | 234 |  |  | -0.67 | 100.5 | 17600.17 | -0.00021 | 0.00135 | -0.021 | 0.0135 | H | neg | neg
156 | PAMF Pre-COVID | Pre-Pandemic | Negative | 235 |  |  | 22 | 85.67 | 19212.17 | 0.00099 | 0.00047 | 0.099 | 0.0047 | H | neg | neg
157 | PAMF Pre-COVID | Pre-Pandemic | Negative | 236 |  |  | -32 | 96 | 18721.83 | -0.00172 | 0.00248 | -0.172 | 0.0248 | H | neg | neg
158 | PAMF Pre-COVID | Pre-Pandemic | Negative | 237 |  |  | 14.67 | 107.33 | 16448.67 | 0.00088 | 0.00352 | 0.088 | 0.0352 | H | neg | neg
159 | PAMF Pre-COVID | Pre-Pandemic | Negative | 238 |  |  | -104.67 | 86 | 13723.33 | -0.00764 | 0.00266 | -0.764 | 0.0266 | H | neg | neg
160 | PAMF Pre-COVID | Pre-Pandemic | Negative | 239 |  |  | -1 | 132.83 | 13820.5 | -8e-05 | 0.00603 | -0.008 | 0.0603 | H | neg | neg
161 | PAMF Pre-COVID | Pre-Pandemic | Negative | 240 |  |  | -6 | 102.83 | 15764.33 | -0.00039 | 0.00338 | -0.039 | 0.0338 | H | neg | neg
162 | PAMF Pre-COVID | Pre-Pandemic | Negative | 241 | HIV |  | -34 | 91.33 | 14834 | -0.0023 | 0.00282 | -0.23 | 0.0282 | H | neg | neg
163 | PAMF Pre-COVID | Pre-Pandemic | Negative | 242 |  |  | -8.33 | 89.17 | 14473.17 | -0.00059 | 0.00274 | -0.059 | 0.0274 | H | neg | neg
164 | PAMF Pre-COVID | Pre-Pandemic | Negative | 243 |  |  | -5.17 | 88 | 16581.83 | -0.00032 | 0.00232 | -0.032 | 0.0232 | H | neg | neg
165 | PAMF Pre-COVID | Pre-Pandemic | Negative | 244 |  |  | 1.17 | 158.67 | 15070.17 | 7e-05 | 0.00724 | 0.007 | 0.0724 | H | neg | neg
166 | PAMF Pre-COVID | Pre-Pandemic | Negative | 245 |  |  | -405 | 104 | 21562.83 | -0.01879 | 0.00253 | -1.879 | 0.0253 | H | neg | neg
167 | PAMF Pre-COVID | Pre-Pandemic | Negative | 247 | HIV |  | 168 | 118.3333 | 12739.8333 | 0.013056 | 0.0034799 | 1.3056 | 0.034799 | H | neg | neg
168 | PAMF Pre-COVID | Pre-Pandemic | Negative | 248 |  |  | -109.5 | 3 | 18952.33 | -0.00592 | -0.00153 | -0.592 | -0.0153 | H | neg | neg
169 | PAMF Pre-COVID | Pre-Pandemic | Negative | 249 |  |  | -97 | 206.33 | 17672.83 | -0.00564 | 0.00986 | -0.564 | 0.0986 | H | neg | neg
170 | PAMF Pre-COVID | Pre-Pandemic | Negative | 250 |  |  | -107 | 85.67 | 15799.33 | -0.00694 | 0.0034 | -0.694 | 0.034 | H | neg | neg
171 | PAMF Pre-COVID | Pre-Pandemic | Negative | 251 |  |  | -113 | 88.67 | 17219.5 | -0.00672 | 0.00329 | -0.672 | 0.0329 | H | neg | neg
172 | PAMF Pre-COVID | Pre-Pandemic | Negative | 252 |  |  | 0.33 | 68.67 | 15872.67 | -0.00015 | 0.00231 | -0.015 | 0.0231 | H | neg | neg
173 | PAMF Pre-COVID | Pre-Pandemic | Negative | 253 |  |  | -3.67 | 97 | 17081.5 | -0.00037 | 0.00381 | -0.037 | 0.0381 | H | neg | neg
174 | PAMF Pre-COVID | Pre-Pandemic | Negative | 254 |  |  | -0.67 | 99 | 19401.33 | -0.00017 | 0.00345 | -0.017 | 0.0345 | H | neg | neg
175 | PAMF Pre-COVID | Pre-Pandemic | Negative | 255 |  |  | 4.83 | 102.5 | 19562.5 | 0.00011 | 0.0036 | 0.011 | 0.036 | H | neg | neg
176 | PAMF Pre-COVID | Pre-Pandemic | Negative | 256 |  |  | -1.67 | 85.83 | 17179.33 | -0.00025 | 0.00313 | -0.025 | 0.0313 | H | neg | neg
177 | PAMF Pre-COVID | Pre-Pandemic | Negative | 257 |  |  | 29.33 | 101.33 | 21580.67 | 0.00124 | 0.00321 | 0.124 | 0.0321 | H | neg | neg
178 | PAMF Pre-COVID | Pre-Pandemic | Negative | 258 |  |  | -2 | 92.33 | 17783.67 | -0.00026 | 0.00339 | -0.026 | 0.0339 | H | neg | neg
179 | PAMF Pre-COVID | Pre-Pandemic | Negative | 259 |  |  | 62.83 | 107.33 | 19244.17 | 0.00313 | 0.00391 | 0.313 | 0.0391 | H | neg | neg
180 | PAMF Pre-COVID | Pre-Pandemic | Negative | 260 |  |  | -9.5 | 100.17 | 18253.83 | -0.00076 | 0.00151 | -0.076 | 0.0151 | H | neg | neg
181 | PAMF Pre-COVID | Pre-Pandemic | Negative | 261 |  |  | -65.33 | 47.33 | 15633.17 | -0.00446 | -0.00162 | -0.446 | -0.0162 | H | neg | neg
182 | PAMF Pre-COVID | Pre-Pandemic | Negative | 262 |  |  | -39.33 | 67.17 | 16181.33 | -0.0027 | -0.00034 | -0.27 | -0.0034 | H | neg | neg
183 | PAMF Pre-COVID | Pre-Pandemic | Negative | 263 |  |  | 1.83 | 101.67 | 16241.83 | -0.00015 | 0.00179 | -0.015 | 0.0179 | H | neg | neg
184 | PAMF Pre-COVID | Pre-Pandemic | Negative | 264 |  |  | -1.83 | 87.5 | 16657.17 | -0.00037 | 0.00089 | -0.037 | 0.0089 | H | neg | neg
185 | PAMF Pre-COVID | Pre-Pandemic | Negative | 265 |  |  | 66.33 | 108 | 17490.5 | 0.00354 | 0.00202 | 0.354 | 0.0202 | H | neg | neg
186 | PAMF Pre-COVID | Pre-Pandemic | Negative | 266 |  |  | 4 | 45.17 | 17403.33 | -2e-05 | -0.00158 | -0.002 | -0.0158 | H | neg | neg
187 | PAMF Pre-COVID | Pre-Pandemic | Negative | 267 |  |  | 39.67 | 106.67 | 20756.33 | 0.0017 | 0.00164 | 0.17 | 0.0164 | H | neg | neg
188 | PAMF Pre-COVID | Pre-Pandemic | Negative | 268 |  |  | 107 | 64.33 | 18670.33 | 0.0055 | -0.00045 | 0.55 | -0.0045 | H | neg | neg
189 | PAMF Pre-COVID | Pre-Pandemic | Negative | 269 |  |  | 3.83 | 93 | 22080.17 | -2e-05 | 0.00092 | -0.002 | 0.0092 | H | neg | neg
190 | PAMF Pre-COVID | Pre-Pandemic | Negative | 270 |  |  | 67.67 | 100.67 | 17024.5 | 0.00372 | 0.00164 | 0.372 | 0.0164 | H | neg | neg
191 | PAMF Pre-COVID | Pre-Pandemic | Negative | 271 |  |  | 64.67 | 110.5 | 18109 | 0.00333 | 0.00209 | 0.333 | 0.0209 | H | neg | neg
192 | PAMF Pre-COVID | Pre-Pandemic | Negative | 272 |  |  | -103.33 | 13.33 | 16386.5 | -0.00649 | -0.00393 | -0.649 | -0.0393 | H | neg | neg
193 | PAMF Pre-COVID | Pre-Pandemic | Negative | 273 |  |  | -58 | 13.67 | 17368.17 | -0.00351 | -0.00368 | -0.351 | -0.0368 | H | neg | neg
194 | PAMF Pre-COVID | Pre-Pandemic | Negative | 274 |  |  | 6.33 | 94.17 | 16833 | 0.0002 | 0.00098 | 0.02 | 0.0098 | H | neg | neg
195 | PAMF Pre-COVID | Pre-Pandemic | Negative | 275 |  |  | -2.33 | 96.83 | 17430.5 | -0.00031 | 0.0011 | -0.031 | 0.011 | H | neg | neg
196 | PAMF Pre-COVID | Pre-Pandemic | Negative | 276 |  |  | -2.67 | 110.17 | 17963 | -0.00032 | 0.00181 | -0.032 | 0.0181 | H | neg | neg
197 | PAMF Pre-COVID | Pre-Pandemic | Negative | 277 |  |  | 3.33 | 102.5 | 18504.67 | 2e-05 | 0.00134 | 0.002 | 0.0134 | H | neg | neg
198 | PAMF Pre-COVID | Pre-Pandemic | Negative | 278 |  |  | -2.83 | 13 | 20290.17 | -0.00029 | -0.00319 | -0.029 | -0.0319 | H | neg | neg
199 | PAMF Pre-COVID | Pre-Pandemic | Negative | 279 |  |  | 59.33 | 72 | 20903.33 | 0.00269 | -0.00027 | 0.269 | -0.0027 | H | neg | neg
200 | PAMF Pre-COVID | Pre-Pandemic | Negative | 280 |  |  | 3.67 | 15 | 19366.33 | 3e-05 | -0.00324 | 0.003 | -0.0324 | H | neg | neg
201 | PAMF Pre-COVID | Pre-Pandemic | Negative | 281 |  |  | 56.33 | 399.17 | 21866.67 | 0.00244 | 0.0147 | 0.244 | 0.147 | H | neg | neg
202 | PAMF Pre-COVID | Pre-Pandemic | Negative | 282 |  |  | 43 | 107.17 | 18471.83 | 0.00217 | 0.0016 | 0.217 | 0.016 | H | neg | neg
203 | PAMF Pre-COVID | Pre-Pandemic | Negative | 283 |  |  | 33.33 | 78 | 21685 | 0.0014 | 2e-05 | 0.14 | 0.0002 | H | neg | neg
204 | PAMF Pre-COVID | Pre-Pandemic | Negative | 284 |  |  | -98 | 67.5 | 17306.67 | -0.0063 | -0.00065 | -0.63 | -0.0065 | H | neg | neg
205 | PAMF Pre-COVID | Pre-Pandemic | Negative | 285 |  |  | -3.17 | 9.67 | 15996.5 | -0.00089 | -0.00431 | -0.089 | -0.0431 | H | neg | neg
206 | PAMF Pre-COVID | Pre-Pandemic | Negative | 286 |  |  | -3.67 | 4.67 | 16829.17 | -0.00087 | -0.0044 | -0.087 | -0.044 | H | neg | neg
207 | PAMF Pre-COVID | Pre-Pandemic | Negative | 287 |  |  | 0.83 | 110.33 | 17600.67 | -0.00058 | 0.0018 | -0.058 | 0.018 | H | neg | neg
208 | PAMF Pre-COVID | Pre-Pandemic | Negative | 288 |  |  | 2 | 43 | 18249.17 | -0.00049 | -0.00195 | -0.049 | -0.0195 | H | neg | neg
209 | PAMF Pre-COVID | Pre-Pandemic | Negative | 289 |  |  | 9.67 | 138.5 | 18586.33 | -7e-05 | 0.00322 | -0.007 | 0.0322 | H | neg | neg
210 | PAMF Pre-COVID | Pre-Pandemic | Negative | 290 |  |  | -2 | 110.67 | 19234.33 | -0.00068 | 0.00166 | -0.068 | 0.0166 | H | neg | neg
211 | PAMF Pre-COVID | Pre-Pandemic | Negative | 291 |  |  | -2 | 81.67 | 22343.17 | -0.00058 | 0.00013 | -0.058 | 0.0013 | H | neg | neg
212 | PAMF Pre-COVID | Pre-Pandemic | Negative | 292 |  |  | -0.67 | 78.33 | 19945.67 | -0.00058 | -2e-05 | -0.058 | -0.0002 | H | neg | neg
213 | PAMF Pre-COVID | Pre-Pandemic | Negative | 293 |  |  | -0.83 | 82.5 | 23204.33 | -0.00051 | 0.00017 | -0.051 | 0.0017 | H | neg | neg
214 | PAMF Pre-COVID | Pre-Pandemic | Negative | 294 |  |  | -33.5 | 102 | 17095.83 | -0.0026 | 0.00136 | -0.26 | 0.0136 | H | neg | neg
215 | PAMF Pre-COVID | Pre-Pandemic | Negative | 295 |  |  | 5 | 121.5 | 24044.33 | -0.00025 | 0.00178 | -0.025 | 0.0178 | H | neg | neg
216 | PAMF Pre-COVID | Pre-Pandemic | Negative | 382 |  |  | 184.5 | 270.3333 | 13430 | 0.013614 | 0.014619 | 1.3614 | 0.14619 | H | neg | neg
217 | PAMF Pre-COVID | Pre-Pandemic | Negative | 384 |  |  | 67.6667 | 230.3333 | 13348.8333 | 0.0049443 | 0.011711 | 0.49443 | 0.11711 | H | neg | neg
218 | PAMF Pre-COVID | Pre-Pandemic | Negative | 390 |  |  | 0.33333 | 92.3333 | 22338 | -0.0030814 | 0.0011042 | -0.30814 | 0.011042 | H | neg | neg
219 | PAMF Pre-COVID | Pre-Pandemic | Negative | 391 |  |  | 125 | 112.3333 | 13523 | 0.00668 | 0.004289 | 0.668 | 0.04289 | H | neg | neg
220 | PAMF Pre-COVID | Pre-Pandemic | Negative | 392 |  |  | 137.5 | 78 | 13608.1667 | 0.0075567 | 0.0017392 | 0.75567 | 0.017392 | H | neg | neg
221 | PAMF Pre-COVID | Pre-Pandemic | Negative | 393 |  |  | 70.6667 | 106.6667 | 13685.6667 | 0.0026305 | 0.003824 | 0.26305 | 0.03824 | H | neg | neg
222 | PAMF Pre-COVID | Pre-Pandemic | Negative | 394 |  |  | 73.3333 | 401.1667 | 11863.5 | 0.0032593 | 0.029235 | 0.32593 | 0.29235 | H | neg | neg
223 | PAMF Pre-COVID | Pre-Pandemic | Negative | 395 |  |  | 205.3333 | 104.6667 | 13438.6667 | 0.0127 | 0.0037454 | 1.27 | 0.037454 | H | neg | neg
224 | PAMF Pre-COVID | Pre-Pandemic | Negative | 396 |  |  | 90.3333 | 104.6667 | 12232.6667 | 0.0045507 | 0.0041147 | 0.45507 | 0.041147 | H | neg | neg
225 | PAMF Pre-COVID | Pre-Pandemic | Negative | 397 |  |  | 97.5 | 68.3333 | 13940.8333 | 0.0045071 | 0.0010042 | 0.45071 | 0.010042 | H | neg | neg
226 | PAMF Pre-COVID | Pre-Pandemic | Negative | 398 |  |  | 155.1667 | 110.3333 | 12392.6667 | 0.0097235 | 0.0045188 | 0.97235 | 0.045188 | H | neg | neg
227 | PAMF Pre-COVID | Pre-Pandemic | Negative | 399 |  |  | 73.5 | 134 | 12691.6667 | 0.0030598 | 0.0062771 | 0.30598 | 0.062771 | H | neg | neg
228 | PAMF Pre-COVID | Pre-Pandemic | Negative | 400 |  |  | -43.1667 | 103 | 17002.6667 | -0.0022644 | 0.0051953 | -0.22644 | 0.051953 | H | neg | neg
229 | PAMF Pre-COVID | Pre-Pandemic | Negative | 401 |  |  | -202.3333 | 86 | 15399.8333 | -0.012836 | 0.0046321 | -1.2836 | 0.046321 | H | neg | neg
230 | PAMF Pre-COVID | Pre-Pandemic | Negative | 402 |  |  | -8.6667 | 58.3333 | 15538.8333 | -0.00025742 | 0.0028102 | -0.025742 | 0.028102 | H | neg | neg
231 | PAMF Pre-COVID | Pre-Pandemic | Negative | 403 |  |  | 1.6667 | 85.6667 | 13433.8333 | 0.00047145 | 0.0052852 | 0.047145 | 0.052852 | H | neg | neg
232 | PAMF Pre-COVID | Pre-Pandemic | Negative | 404 |  |  | -13.6667 | 159 | 14607 | -0.00061614 | 0.0098811 | -0.061614 | 0.098811 | H | neg | neg
233 | PAMF Pre-COVID | Pre-Pandemic | Negative | 405 |  |  | 0.83333 | 109 | 14789.5 | 0.00037189 | 0.0063784 | 0.037189 | 0.063784 | H | neg | neg
234 | PAMF Pre-COVID | Pre-Pandemic | Negative | 406 |  |  | -4.6667 | 97.3333 | 12936.3333 | 0 | 0.0063903 | 0 | 0.063903 | H | neg | neg
235 | PAMF Pre-COVID | Pre-Pandemic | Negative | 407 |  |  | 58 | 89.3333 | 13796.3333 | 0.0045423 | 0.0054121 | 0.45423 | 0.054121 | H | neg | neg
236 | PAMF Pre-COVID | Pre-Pandemic | Negative | 408 |  |  | 39.3333 | 91 | 12456.1667 | 0.0035324 | 0.0061282 | 0.35324 | 0.061282 | H | neg | neg
237 | PAMF Pre-COVID | Pre-Pandemic | Negative | 409 |  |  | 10.1667 | 80.6667 | 12993 | 0.0011416 | 0.0050797 | 0.11416 | 0.050797 | H | neg | neg
238 | PAMF Pre-COVID | Pre-Pandemic | Negative | 410 |  |  | 0.66667 | 92.3333 | 12688 | 0.00042034 | 0.0061213 | 0.042034 | 0.061213 | H | neg | neg
239 | PAMF Pre-COVID | Pre-Pandemic | Negative | 411 |  |  | 63.5 | 90.8333 | 13296.3333 | 0.0051267 | 0.0057284 | 0.51267 | 0.057284 | H | neg | neg
240 | PAMF Pre-COVID | Pre-Pandemic | Negative | 412 |  |  | -105.3333 | 83.6667 | 17555.5 | -0.0063323 | 0.0006076 | -0.63323 | 0.006076 | H | neg | neg
241 | PAMF Pre-COVID | Pre-Pandemic | Negative | 413 |  |  | -34.6667 | 12.8333 | 18262 | -0.0022177 | -0.0032946 | -0.22177 | -0.032946 | H | neg | neg
242 | PAMF Pre-COVID | Pre-Pandemic | Negative | 414 |  |  | 5.3333 | 83.3333 | 16007.1667 | -3.1236e-05 | 0.00064554 | -0.0031236 | 0.0064554 | H | neg | neg
243 | PAMF Pre-COVID | Pre-Pandemic | Negative | 415 |  |  | -5.6667 | 76 | 16834 | -0.00068314 | 0.00017821 | -0.068314 | 0.0017821 | H | neg | neg
244 | PAMF Pre-COVID | Pre-Pandemic | Negative | 416 |  |  | -35.3333 | 40.6667 | 16613.8333 | -0.0024779 | -0.0019462 | -0.24779 | -0.019462 | H | neg | neg
245 | PAMF Pre-COVID | Pre-Pandemic | Negative | 417 |  |  | 0 | 184.3333 | 16928.6667 | -0.00034458 | 0.0065766 | -0.034458 | 0.065766 | H | neg | neg
246 | PAMF Pre-COVID | Pre-Pandemic | Negative | 418 |  |  | 11 | 90.6667 | 14634.3333 | 0.00035305 | 0.0012072 | 0.035305 | 0.012072 | H | neg | neg
247 | PAMF Pre-COVID | Pre-Pandemic | Negative | 419 |  |  | -0.16667 | 1747.1667 | 15465.1667 | -0.00038797 | 0.10825 | -0.038797 | 1.0825 | H | neg | neg
248 | PAMF Pre-COVID | Pre-Pandemic | Negative | 420 |  |  | -6.6667 | 91.3333 | 13739 | -0.00090982 | 0.0013344 | -0.090982 | 0.013344 | H | neg | neg
249 | PAMF Pre-COVID | Pre-Pandemic | Negative | 421 |  |  | -7.8333 | 119.3333 | 16254.1667 | -0.00084081 | 0.0028506 | -0.084081 | 0.028506 | H | neg | neg
250 | PAMF Pre-COVID | Pre-Pandemic | Negative | 422 |  |  | 1.8333 | 57.3333 | 13990.1667 | -0.00028592 | -0.0011198 | -0.028592 | -0.011198 | H | neg | neg
251 | PAMF Pre-COVID | Pre-Pandemic | Negative | 423 |  |  | 89.6667 | 126 | 16117.8333 | 0.0052013 | 0.0032883 | 0.52013 | 0.032883 | H | neg | neg
252 | PAMF Pre-COVID | Pre-Pandemic | Negative | 424 |  |  | -69.6667 | 121 | 18390.6667 | -0.0054738 | 0.0025647 | -0.54738 | 0.025647 | H | neg | neg
253 | PAMF Pre-COVID | Pre-Pandemic | Negative | 425 |  |  | -32.3333 | 101 | 18836.1667 | -0.0033623 | 0.0014423 | -0.33623 | 0.014423 | H | neg | neg
254 | PAMF Pre-COVID | Pre-Pandemic | Negative | 426 |  |  | 0 | 81.3333 | 17348.1667 | -0.0017869 | 0.00043232 | -0.17869 | 0.0043232 | H | neg | neg
255 | PAMF Pre-COVID | Pre-Pandemic | Negative | 427 |  |  | -0.5 | 156 | 18270.8333 | -0.0017241 | 0.0044971 | -0.17241 | 0.044971 | H | neg | neg
256 | PAMF Pre-COVID | Pre-Pandemic | Negative | 428 |  |  | -6.1667 | 101.6667 | 19137.8333 | -0.0019421 | 0.0014544 | -0.19421 | 0.014544 | H | neg | neg
257 | PAMF Pre-COVID | Pre-Pandemic | Negative | 429 |  |  | -0.83333 | 88 | 18800.3333 | -0.0016932 | 0.00075353 | -0.16932 | 0.0075353 | H | neg | neg
258 | PAMF Pre-COVID | Pre-Pandemic | Negative | 430 |  |  | 1 | 157.8333 | 17776.6667 | -0.0016876 | 0.0047253 | -0.16876 | 0.047253 | H | neg | neg
259 | PAMF Pre-COVID | Pre-Pandemic | Negative | 431 |  |  | 103.6667 | 196.8333 | 16578.5 | 0.0043832 | 0.0074192 | 0.43832 | 0.074192 | H | neg | neg
260 | PAMF Pre-COVID | Pre-Pandemic | Negative | 432 |  |  | 109.3333 | 96.5 | 16218.3333 | 0.0048299 | 0.0013976 | 0.48299 | 0.013976 | H | neg | neg
261 | PAMF Pre-COVID | Pre-Pandemic | Negative | 433 |  |  | 100 | 17.5 | 17573.1667 | 0.0039264 | -0.0032056 | 0.39264 | -0.032056 | H | neg | neg
262 | PAMF Pre-COVID | Pre-Pandemic | Negative | 434 |  |  | 279.5 | 3180.5 | 13540.3333 | 0.020519 | 0.22943 | 2.0519 | 2.2943 | H | pos | pos
263 | PAMF Pre-COVID | Pre-Pandemic | Negative | 435 |  |  | 4 | 64 | 16584.5 | -0.001628 | -0.00059292 | -0.1628 | -0.0059292 | H | neg | neg
264 | PAMF Pre-COVID | Pre-Pandemic | Negative | 436 |  |  | 104 | 359.6667 | 12917.1667 | 0.0079223 | 0.022115 | 0.79223 | 0.22115 | H | neg | neg
265 | PAMF Pre-COVID | Pre-Pandemic | Negative | 437 |  |  | -9 | 86.6667 | 15454.8333 | -0.001251 | 0.00069018 | -0.1251 | 0.0069018 | H | neg | neg
266 | PAMF Pre-COVID | Pre-Pandemic | Negative | 438 |  |  | -104 | 84.3333 | 15737.8333 | -0.0072649 | 0.00052951 | -0.72649 | 0.0052951 | H | neg | neg
267 | PAMF Pre-COVID | Pre-Pandemic | Negative | 439 |  |  | -82.3333 | 40.8333 | 16517 | -0.0056104 | -0.0021291 | -0.56104 | -0.021291 | H | neg | neg
268 | PAMF Pre-COVID | Pre-Pandemic | Negative | 440 |  |  | -6.1667 | 102.3333 | 17890.6667 | -0.00092227 | 0.0014719 | -0.092227 | 0.014719 | H | neg | neg
269 | PAMF Pre-COVID | Pre-Pandemic | Negative | 441 |  |  | -34.1667 | 80.6667 | 19367.6667 | -0.0022976 | 0.00024095 | -0.22976 | 0.0024095 | H | neg | neg
270 | PAMF Pre-COVID | Pre-Pandemic | Negative | 442 |  |  | -6 | 101.3333 | 18910.8333 | -0.0008637 | 0.0013396 | -0.08637 | 0.013396 | H | neg | neg
271 | PAMF Pre-COVID | Pre-Pandemic | Negative | 443 |  |  | 12 | 97.5 | 18771.1667 | 8.8789e-05 | 0.0011454 | 0.0088789 | 0.011454 | H | neg | neg
272 | PAMF Pre-COVID | Pre-Pandemic | Negative | 444 |  |  | -2.6667 | 90.3333 | 19214.1667 | -0.00067658 | 0.00074598 | -0.067658 | 0.0074598 | H | neg | neg
273 | PAMF Pre-COVID | Pre-Pandemic | Negative | 445 | HIV |  | 12.1667 | 185.8333 | 20510.3333 | 8.9386e-05 | 0.005355 | 0.0089386 | 0.05355 | H | neg | neg
274 | PAMF Pre-COVID | Pre-Pandemic | Negative | 446 |  |  | 102.6667 | 89.3333 | 19281.1667 | 0.0047888 | 0.00069152 | 0.47888 | 0.0069152 | H | neg | neg
275 | PAMF Pre-COVID | Pre-Pandemic | Negative | 447 |  |  | -0.33333 | 125.6667 | 19724.1667 | -0.00054079 | 0.0025181 | -0.054079 | 0.025181 | H | neg | neg
276 | PAMF Pre-COVID | Pre-Pandemic | Negative | 448 |  |  | 6.1667 | 60.6667 | 27832.3333 | -0.0022695 | -0.00037127 | -0.22695 | -0.0037127 | H | neg | neg
277 | PAMF Pre-COVID | Pre-Pandemic | Negative | 449 |  |  | 90.3333 | 89.1667 | 26702.6667 | 0.00078644 | 0.00068033 | 0.078644 | 0.0068033 | H | neg | neg
278 | PAMF Pre-COVID | Pre-Pandemic | Negative | 450 |  |  | -33.8333 | 68.5 | 24334.6667 | -0.0042395 | -0.00010273 | -0.42395 | -0.0010273 | H | neg | neg
279 | PAMF Pre-COVID | Pre-Pandemic | Negative | 451 |  |  | -2 | 15 | 22840.3333 | -0.0031231 | -0.0024518 | -0.31231 | -0.024518 | H | neg | neg
280 | PAMF Pre-COVID | Pre-Pandemic | Negative | 452 |  |  | 139.1667 | 67 | 23203.1667 | 0.0030096 | -0.00017239 | 0.30096 | -0.0017239 | H | neg | neg
281 | PAMF Pre-COVID | Pre-Pandemic | Negative | 453 |  |  | 58.3333 | 90 | 21597.5 | -0.00050932 | 0.00087973 | -0.050932 | 0.0087973 | H | neg | neg
282 | PAMF Pre-COVID | Pre-Pandemic | Negative | 454 |  |  | -1.3333 | 177.3333 | 22008.6667 | -0.0032109 | 0.0048314 | -0.32109 | 0.048314 | H | neg | neg
283 | PAMF Pre-COVID | Pre-Pandemic | Negative | 455 |  |  | 94.3333 | 91.5 | 21536.8333 | 0.0011608 | 0.00095186 | 0.11608 | 0.0095186 | H | neg | neg
284 | PAMF Pre-COVID | Pre-Pandemic | Negative | 456 |  |  | 3 | 93 | 23460.6667 | -0.0028274 | 0.00093774 | -0.28274 | 0.0093774 | H | neg | neg
285 | PAMF Pre-COVID | Pre-Pandemic | Negative | 457 |  |  | 25.8333 | 91.5 | 21935.6667 | -0.0019831 | 0.00093455 | -0.19831 | 0.0093455 | H | neg | neg
286 | PAMF Pre-COVID | Pre-Pandemic | Negative | 458 |  |  | 107.1667 | 275.5 | 21993.3333 | 0.0017202 | 0.0092983 | 0.17202 | 0.092983 | H | neg | neg
287 | PAMF Pre-COVID | Pre-Pandemic | Negative | 459 |  |  | 105.3333 | 197.5 | 22242 | 0.0016186 | 0.0056874 | 0.16186 | 0.056874 | H | neg | neg
288 | PAMF Pre-COVID | Pre-Pandemic | Negative | 460 |  |  | 89.6667 | 93.6667 | 26604.3333 | 0.0022678 | 0.00055129 | 0.22678 | 0.0055129 | H | neg | neg
289 | PAMF Pre-COVID | Pre-Pandemic | Negative | 461 |  |  | 103.6667 | 103.3333 | 26880 | 0.0027654 | 0.00090526 | 0.27654 | 0.0090526 | H | neg | neg
290 | PAMF Pre-COVID | Pre-Pandemic | Negative | 462 |  |  | 0 | 93.3333 | 24194.8333 | -0.0012124 | 0.00059241 | -0.12124 | 0.0059241 | H | neg | neg
291 | PAMF Pre-COVID | Pre-Pandemic | Negative | 463 |  |  | -2.3333 | 95.3333 | 23432.6667 | -0.0013514 | 0.00069703 | -0.13514 | 0.0069703 | H | neg | neg
292 | PAMF Pre-COVID | Pre-Pandemic | Negative | 464 |  |  | 97 | 97.6667 | 23576.6667 | 0.0028701 | 0.00079174 | 0.28701 | 0.0079174 | H | neg | neg
293 | PAMF Pre-COVID | Pre-Pandemic | Negative | 465 |  |  | 94.3333 | 155.3333 | 21593 | 0.0030102 | 0.0035351 | 0.30102 | 0.035351 | H | neg | neg
294 | PAMF Pre-COVID | Pre-Pandemic | Negative | 466 |  |  | 102.6667 | 219.6667 | 23017 | 0.0031861 | 0.0061114 | 0.31861 | 0.061114 | H | neg | neg
295 | PAMF Pre-COVID | Pre-Pandemic | Negative | 467 |  |  | 72 | 67.3333 | 18886.8333 | 0.0022591 | -0.00061771 | 0.22591 | -0.0061771 | H | neg | neg
296 | PAMF Pre-COVID | Pre-Pandemic | Negative | 468 |  |  | 101.3333 | 156.5 | 22710.6667 | 0.0031703 | 0.0034125 | 0.31703 | 0.034125 | H | neg | neg
297 | PAMF Pre-COVID | Pre-Pandemic | Negative | 469 |  |  | 30 | 92.3333 | 19578.3333 | 3.4051e-05 | 0.00068102 | 0.0034051 | 0.0068102 | H | neg | neg
298 | PAMF Pre-COVID | Pre-Pandemic | Negative | 470 |  |  | 117 | 110.1667 | 20032.8333 | 0.0043761 | 0.0015558 | 0.43761 | 0.015558 | H | neg | neg
299 | PAMF Pre-COVID | Pre-Pandemic | Negative | 471 |  |  | 93 | 93.3333 | 20123.3333 | 0.0031638 | 0.00071227 | 0.31638 | 0.0071227 | H | neg | neg
300 | PAMF Pre-COVID | Pre-Pandemic | Negative | 472 |  |  | -280 | 101.6667 | 27230.8333 | -0.013906 | 0.00078343 | -1.3906 | 0.0078343 | H | neg | neg
301 | PAMF Pre-COVID | Pre-Pandemic | Negative | 473 |  |  | 56.6667 | 95.3333 | 29140.1667 | -0.0014413 | 0.00051475 | -0.14413 | 0.0051475 | H | neg | neg
302 | PAMF Pre-COVID | Pre-Pandemic | Negative | 474 |  |  | -4.5 | 94 | 26540 | -0.0038872 | 0.00051495 | -0.38872 | 0.0051495 | H | neg | neg
303 | PAMF Pre-COVID | Pre-Pandemic | Negative | 475 |  |  | 4 | 33.6667 | 25150.3333 | -0.003764 | -0.0018555 | -0.3764 | -0.018555 | H | neg | neg
304 | PAMF Pre-COVID | Pre-Pandemic | Negative | 476 |  |  | 37.5 | 204.6667 | 26171 | -0.0023372 | 0.0047508 | -0.23372 | 0.047508 | H | neg | neg
305 | PAMF Pre-COVID | Pre-Pandemic | Negative | 477 |  |  | 2 | 93.3333 | 24194 | -0.0039955 | 0.00053732 | -0.39955 | 0.0053732 | H | neg | neg
306 | PAMF Pre-COVID | Pre-Pandemic | Negative | 478 |  |  | 95.3333 | 191.6667 | 26561.5 | -0.00012549 | 0.0041915 | -0.012549 | 0.041915 | H | neg | neg
307 | PAMF Pre-COVID | Pre-Pandemic | Negative | 479 |  |  | 69.3333 | 105.1667 | 24794 | -0.0011831 | 0.0010016 | -0.11831 | 0.010016 | H | neg | neg
308 | PAMF Pre-COVID | Pre-Pandemic | Negative | 480 |  |  | 94.8333 | 108.1667 | 25701.1667 | -0.00014915 | 0.001083 | -0.014915 | 0.01083 | H | neg | neg
309 | PAMF Pre-COVID | Pre-Pandemic | Negative | 481 |  |  | 68 | 92 | 25728.3333 | -0.0011919 | 0.00045346 | -0.11919 | 0.0045346 | H | neg | neg
310 | PAMF Pre-COVID | Pre-Pandemic | Negative | 482 |  |  | 97.6667 | 97 | 22704 | -4.4045e-05 | 0.00073409 | -0.0044045 | 0.0073409 | H | neg | neg
311 | PAMF Pre-COVID | Pre-Pandemic | Negative | 483 |  |  | 8.3333 | 105.6667 | 23337.8333 | -0.0038707 | 0.0010855 | -0.38707 | 0.010855 | H | neg | neg
312 | PAMF Pre-COVID | Pre-Pandemic | Negative | 484 |  |  | -36 | 87.6667 | 31071.5 | -0.0012123 | 0.00030038 | -0.12123 | 0.0030038 | H | neg | neg
313 | PAMF Pre-COVID | Pre-Pandemic | Negative | 485 |  |  | -0.83333 | 90.8333 | 33278.3333 | -7.5124e-05 | 0.00037562 | -0.0075124 | 0.0037562 | H | neg | neg
314 | PAMF Pre-COVID | Pre-Pandemic | Negative | 486 |  |  | -7 | 85.6667 | 27169.1667 | -0.00031899 | 0.00026991 | -0.031899 | 0.0026991 | H | neg | neg
315 | PAMF Pre-COVID | Pre-Pandemic | Negative | 487 | HIV |  | 104.6667 | 291 | 13170.3333 | 0.0078206 | 0.016476 | 0.78206 | 0.16476 | H | neg | neg
316 | PAMF Pre-COVID | Pre-Pandemic | Negative | 488 |  |  | 5 | 177.1667 | 27447 | 0.00012145 | 0.0036009 | 0.012145 | 0.036009 | H | neg | neg
317 | PAMF Pre-COVID | Pre-Pandemic | Negative | 489 |  |  | 31 | 104.5 | 29301.5 | 0.0010011 | 0.00089301 | 0.10011 | 0.0089301 | H | neg | neg
318 | PAMF Pre-COVID | Pre-Pandemic | Negative | 490 |  |  | -1.3333 | 109.6667 | 27390.6667 | -0.00010953 | 0.0011439 | -0.010953 | 0.011439 | H | neg | neg
319 | PAMF Pre-COVID | Pre-Pandemic | Negative | 491 |  |  | 27.3333 | 97.3333 | 27159.6667 | 0.00094503 | 0.00069957 | 0.094503 | 0.0069957 | H | neg | neg
320 | PAMF Pre-COVID | Pre-Pandemic | Negative | 492 |  |  | 59.5 | 96 | 26178.8333 | 0.0022092 | 0.00067485 | 0.22092 | 0.0067485 | H | neg | neg
321 | PAMF Pre-COVID | Pre-Pandemic | Negative | 493 | HIV |  | 58.1667 | 91.3333 | 27881.3333 | 0.0020264 | 0.00046626 | 0.20264 | 0.0046626 | H | neg | neg
322 | PAMF Pre-COVID | Pre-Pandemic | Negative | 494 |  |  | 66.6667 | 101.3333 | 24358.1667 | 0.0026685 | 0.00094424 | 0.26685 | 0.0094424 | H | neg | neg
323 | PAMF Pre-COVID | Pre-Pandemic | Negative | 495 |  |  | 34 | 90.3333 | 25766.5 | 0.0012549 | 0.00046572 | 0.12549 | 0.0046572 | H | neg | neg
324 | PAMF Pre-COVID | Pre-Pandemic | Negative | 496 |  |  | -60 | 69 | 21773.6667 | -0.0028322 | 0.0024035 | -0.28322 | 0.024035 | H | neg | neg
325 | PAMF Pre-COVID | Pre-Pandemic | Negative | 497 |  |  | -5.1667 | 40.6667 | 19059.5 | -0.00035853 | 0.0012592 | -0.035853 | 0.012592 | H | neg | neg
326 | PAMF Pre-COVID | Pre-Pandemic | Negative | 498 |  |  | -40 | 94.1667 | 18153.1667 | -0.0022953 | 0.0042692 | -0.22953 | 0.042692 | H | neg | neg
327 | PAMF Pre-COVID | Pre-Pandemic | Negative | 499 |  |  | 1 | 60.3333 | 16761.8333 | -3.9773e-05 | 0.0026051 | -0.0039773 | 0.026051 | H | neg | neg
328 | PAMF Pre-COVID | Pre-Pandemic | Negative | 500 |  |  | -2.1667 | 69.3333 | 18523.3333 | -0.00020695 | 0.0028433 | -0.020695 | 0.028433 | H | neg | neg
329 | PAMF Pre-COVID | Pre-Pandemic | Negative | 501 |  |  | 1.3333 | 83.3333 | 16648.3333 | -2.0022e-05 | 0.0040044 | -0.0020022 | 0.040044 | H | neg | neg
330 | PAMF Pre-COVID | Pre-Pandemic | Negative | 502 |  |  | 67.6667 | 88.5 | 17442.6667 | 0.0037838 | 0.0041183 | 0.37838 | 0.041183 | H | neg | neg
331 | PAMF Pre-COVID | Pre-Pandemic | Negative | 503 |  |  | 0.83333 | 91 | 16846.8333 | -4.9465e-05 | 0.0044123 | -0.0049465 | 0.044123 | H | neg | neg
332 | PAMF Pre-COVID | Pre-Pandemic | Negative | 504 |  |  | 99.3333 | 84.3333 | 16312 | 0.0059874 | 0.0041483 | 0.59874 | 0.041483 | H | neg | neg
333 | PAMF Pre-COVID | Pre-Pandemic | Negative | 505 |  |  | 95.6667 | 103 | 18434.3333 | 0.0050992 | 0.0046833 | 0.50992 | 0.046833 | H | neg | neg
334 | PAMF Pre-COVID | Pre-Pandemic | Negative | 506 |  |  | -4.6667 | 108 | 17002 | -0.00037251 | 0.0053719 | -0.037251 | 0.053719 | H | neg | neg
335 | PAMF Pre-COVID | Pre-Pandemic | Negative | 507 |  |  | 68.6667 | 59.6667 | 16403.8333 | 0.0040844 | 0.0026213 | 0.40844 | 0.026213 | H | neg | neg
336 | PAMF Pre-COVID | Pre-Pandemic | Negative | 508 |  |  | -32 | 95.3333 | 23029.6667 | -0.0032567 | 0.00098424 | -0.32567 | 0.0098424 | H | neg | neg
337 | PAMF Pre-COVID | Pre-Pandemic | Negative | 509 |  |  | 3.5 | 88 | 22634.6667 | -0.0017451 | 0.00067743 | -0.17451 | 0.0067743 | H | neg | neg
338 | PAMF Pre-COVID | Pre-Pandemic | Negative | 510 |  |  | -34.5 | 16.1667 | 20495.3333 | -0.0037813 | -0.0027567 | -0.37813 | -0.027567 | H | neg | neg
339 | PAMF Pre-COVID | Pre-Pandemic | Negative | 511 |  |  | 3 | 99.3333 | 21285.8333 | -0.0018792 | 0.0012528 | -0.18792 | 0.012528 | H | neg | neg
340 | PAMF Pre-COVID | Pre-Pandemic | Negative | 512 |  |  | 100.6667 | 103 | 19458.6667 | 0.0029635 | 0.0015589 | 0.29635 | 0.015589 | H | neg | neg
341 | PAMF Pre-COVID | Pre-Pandemic | Negative | 513 |  |  | -3 | 86.3333 | 19924.5 | -0.0023087 | 0.00068592 | -0.23087 | 0.0068592 | H | neg | neg
342 | PAMF Pre-COVID | Pre-Pandemic | Negative | 514 |  |  | -66 | 67 | 19055.5 | -0.0057201 | -0.00029738 | -0.57201 | -0.0029738 | H | neg | neg
343 | PAMF Pre-COVID | Pre-Pandemic | Negative | 515 |  |  | 64.6667 | 90.3333 | 18173.5 | 0.0011922 | 0.00097211 | 0.11922 | 0.0097211 | H | neg | neg
344 | PAMF Pre-COVID | Pre-Pandemic | Negative | 516 |  |  | 7.6667 | 69 | 18646.5 | -0.0018949 | -0.00019664 | -0.18949 | -0.0019664 | H | neg | neg
345 | PAMF Pre-COVID | Pre-Pandemic | Negative | 517 |  |  | 72 | 96.3333 | 17635.1667 | 0.0016444 | 0.001342 | 0.16444 | 0.01342 | H | neg | neg
346 | PAMF Pre-COVID | Pre-Pandemic | Negative | 518 |  |  | -2.6667 | 121.5 | 17934.1667 | -0.0025464 | 0.0027229 | -0.25464 | 0.027229 | H | neg | neg
347 | PAMF Pre-COVID | Pre-Pandemic | Negative | 519 |  |  | 97.5 | 37.5 | 16707 | 0.0032621 | -0.0021049 | 0.32621 | -0.021049 | H | neg | neg
348 | PAMF Pre-COVID | Pre-Pandemic | Negative | 520 |  |  | 8.3333 | 62.5 | 27482.8333 | 0.00023045 | -0.00046696 | 0.023045 | -0.0046696 | H | neg | neg
349 | PAMF Pre-COVID | Pre-Pandemic | Negative | 521 |  |  | 105.3333 | 73.3333 | 27242.8333 | 0.003793 | -7.3414e-05 | 0.3793 | -0.00073414 | H | neg | neg
350 | PAMF Pre-COVID | Pre-Pandemic | Negative | 522 |  |  | -3.3333 | 15.1667 | 23286.3333 | -0.00022903 | -0.0025838 | -0.022903 | -0.025838 | H | neg | neg
351 | PAMF Pre-COVID | Pre-Pandemic | Negative | 523 |  |  | 98.3333 | 92.6667 | 22489.8333 | 0.0042834 | 0.00077072 | 0.42834 | 0.0077072 | H | neg | neg
352 | PAMF Pre-COVID | Pre-Pandemic | Negative | 524 |  |  | 2.6667 | 89.6667 | 22142.5 | 3.0108e-05 | 0.00064732 | 0.0030108 | 0.0064732 | H | neg | neg
353 | PAMF Pre-COVID | Pre-Pandemic | Negative | 525 |  |  | -0.33333 | 95 | 21737.1667 | -0.00010734 | 0.00090475 | -0.010734 | 0.0090475 | H | neg | neg
354 | PAMF Pre-COVID | Pre-Pandemic | Negative | 526 |  |  | 105 | 103.8333 | 20544.6667 | 0.0050135 | 0.0013872 | 0.50135 | 0.013872 | H | neg | neg
355 | PAMF Pre-COVID | Pre-Pandemic | Negative | 527 |  |  | 105.6667 | 88.6667 | 19984.6667 | 0.0051873 | 0.00066718 | 0.51873 | 0.0066718 | H | neg | neg
356 | PAMF Pre-COVID | Pre-Pandemic | Negative | 528 |  |  | 103.6667 | 103.3333 | 21841.1667 | 0.0046548 | 0.001282 | 0.46548 | 0.01282 | H | neg | neg
357 | PAMF Pre-COVID | Pre-Pandemic | Negative | 529 |  |  | 195.1667 | 69.1667 | 21978.5 | 0.0087889 | -0.00028058 | 0.87889 | -0.0028058 | H | neg | neg
358 | PAMF Pre-COVID | Pre-Pandemic | Negative | 530 |  |  | 2.6667 | 98.6667 | 21947 | 3.0376e-05 | 0.0010632 | 0.0030376 | 0.010632 | H | neg | neg
359 | PAMF Pre-COVID | Pre-Pandemic | Negative | 531 |  |  | 96 | 95 | 20707.6667 | 0.0045394 | 0.00094973 | 0.45394 | 0.0094973 | H | neg | neg
360 | PAMF Pre-COVID | Pre-Pandemic | Negative | 532 |  |  | 4.3333 | 102 | 30167 | -0.00020994 | 0.00083977 | -0.020994 | 0.0083977 | H | neg | neg
361 | PAMF Pre-COVID | Pre-Pandemic | Negative | 533 |  |  | 65 | 98.5 | 26593 | 0.0020431 | 0.00082102 | 0.20431 | 0.0082102 | H | neg | neg
362 | PAMF Pre-COVID | Pre-Pandemic | Negative | 534 |  |  | -68 | 94.8333 | 25630.1667 | -0.0030693 | 0.0007088 | -0.30693 | 0.007088 | H | neg | neg
363 | PAMF Pre-COVID | Pre-Pandemic | Negative | 535 |  |  | 72.3333 | 56.6667 | 26277.3333 | 0.0023468 | -0.00076111 | 0.23468 | -0.0076111 | H | neg | neg
364 | PAMF Pre-COVID | Pre-Pandemic | Negative | 536 |  |  | 191.6667 | 89 | 25460.8333 | 0.007109 | 0.0004844 | 0.7109 | 0.004844 | H | neg | neg
365 | PAMF Pre-COVID | Pre-Pandemic | Negative | 537 |  |  | -3 | 67 | 24574 | -0.00055614 | -0.00039337 | -0.055614 | -0.0039337 | H | neg | neg
366 | PAMF Pre-COVID | Pre-Pandemic | Negative | 538 |  |  | -11 | 352.8333 | 25328.8333 | -0.00085542 | 0.010903 | -0.085542 | 0.10903 | H | neg | neg
367 | PAMF Pre-COVID | Pre-Pandemic | Negative | 539 |  |  | 88 | 102.8333 | 22913.3333 | 0.003375 | 0.001142 | 0.3375 | 0.01142 | H | neg | neg
368 | PAMF Pre-COVID | Pre-Pandemic | Negative | 540 |  |  | 95.3333 | 130.1667 | 23963.1667 | 0.0035332 | 0.0022326 | 0.35332 | 0.022326 | H | neg | neg
369 | PAMF Pre-COVID | Pre-Pandemic | Negative | 541 |  |  | 75.1667 | 99 | 24881.1667 | 0.0025923 | 0.0008976 | 0.25923 | 0.008976 | H | neg | neg
370 | PAMF Pre-COVID | Pre-Pandemic | Negative | 542 |  |  | 96.3333 | 67 | 24297 | 0.0035258 | -0.00039785 | 0.35258 | -0.0039785 | H | neg | neg
371 | PAMF Pre-COVID | Pre-Pandemic | Negative | 543 |  |  | 97.6667 | 98.6667 | 24258.8333 | 0.0035863 | 0.00090689 | 0.35863 | 0.0090689 | H | neg | neg
372 | PAMF Pre-COVID | Pre-Pandemic | Negative | 544 |  |  | 35 | 79.6667 | 27111 | -0.0014754 | 0.00016598 | -0.14754 | 0.0016598 | H | neg | neg
373 | PAMF Pre-COVID | Pre-Pandemic | Negative | 545 |  |  | 124.3333 | 133.3333 | 24503.3333 | 0.0020133 | 0.0023738 | 0.20133 | 0.023738 | H | neg | neg
374 | PAMF Pre-COVID | Pre-Pandemic | Negative | 546 |  |  | -6.3333 | 88 | 25551 | -0.0031832 | 0.00050226 | -0.31832 | 0.0050226 | H | neg | neg
375 | PAMF Pre-COVID | Pre-Pandemic | Negative | 547 |  |  | 103.6667 | -1.3333 | 21601 | 0.0013271 | -0.0035415 | 0.13271 | -0.035415 | H | neg | neg
376 | PAMF Pre-COVID | Pre-Pandemic | Negative | 548 |  |  | 0.5 | 88.3333 | 21938.8333 | -0.0033958 | 0.00060015 | -0.33958 | 0.0060015 | H | neg | neg
377 | PAMF Pre-COVID | Pre-Pandemic | Negative | 549 |  |  | -2 | 105 | 22547.8333 | -0.003415 | 0.0013231 | -0.3415 | 0.013231 | H | neg | neg
378 | PAMF Pre-COVID | Pre-Pandemic | Negative | 550 |  |  | 40 | 94.3333 | 21491.3333 | -0.0016286 | 0.00089183 | -0.16286 | 0.0089183 | H | neg | neg
379 | PAMF Pre-COVID | Pre-Pandemic | Negative | 551 |  |  | 24.6667 | 100 | 20384 | -0.0024693 | 0.0012183 | -0.24693 | 0.012183 | H | neg | neg
380 | PAMF Pre-COVID | Pre-Pandemic | Negative | 552 |  |  | 90.6667 | 91 | 19873.6667 | 0.00078831 | 0.0007967 | 0.078831 | 0.007967 | H | neg | neg
381 | PAMF Pre-COVID | Pre-Pandemic | Negative | 553 |  |  | 104.6667 | 91.3333 | 20341 | 0.0014585 | 0.00079478 | 0.14585 | 0.0079478 | H | neg | neg
382 | PAMF Pre-COVID | Pre-Pandemic | Negative | 554 |  |  | 104.1667 | 108.8333 | 21483.6667 | 0.0013576 | 0.0015671 | 0.13576 | 0.015671 | H | neg | neg
383 | PAMF Pre-COVID | Pre-Pandemic | Negative | 555 |  |  | 99 | 44 | 18299.5 | 0.0013115 | -0.0017031 | 0.13115 | -0.017031 | H | neg | neg
384 | PAMF Pre-COVID | Pre-Pandemic | Negative | 556 |  |  | 5 | 40.6667 | 24075.8333 | -0.0023675 | -0.0013984 | -0.23675 | -0.013984 | H | neg | neg
385 | PAMF Pre-COVID | Pre-Pandemic | Negative | 557 |  |  | -0.66667 | 31.6667 | 22625.6667 | -0.0027697 | -0.0018858 | -0.27697 | -0.018858 | H | neg | neg
386 | PAMF Pre-COVID | Pre-Pandemic | Negative | 558 |  |  | 9 | 100.3333 | 21224.6667 | -0.0024971 | 0.001225 | -0.24971 | 0.01225 | H | neg | neg
387 | PAMF Pre-COVID | Pre-Pandemic | Negative | 559 |  |  | 72.5 | 251.3333 | 22985.3333 | 0.00045681 | 0.0077006 | 0.045681 | 0.077006 | H | neg | neg
388 | PAMF Pre-COVID | Pre-Pandemic | Negative | 560 |  |  | 83.5 | 102.1667 | 21236.1667 | 0.0010124 | 0.0013107 | 0.10124 | 0.013107 | H | neg | neg
389 | PAMF Pre-COVID | Pre-Pandemic | Negative | 561 |  |  | -0.33333 | 95.1667 | 20875.1667 | -0.002986 | 0.000998 | -0.2986 | 0.00998 | H | neg | neg
390 | PAMF Pre-COVID | Pre-Pandemic | Negative | 562 |  |  | 103 | 195.3333 | 22519.5 | 0.0018206 | 0.0053731 | 0.18206 | 0.053731 | H | neg | neg
391 | PAMF Pre-COVID | Pre-Pandemic | Negative | 563 |  |  | 96.8333 | 111.3333 | 22749.5 | 0.0015312 | 0.0016264 | 0.15312 | 0.016264 | H | neg | neg
392 | PAMF Pre-COVID | Pre-Pandemic | Negative | 564 |  |  | 2.1667 | 87 | 21964.1667 | -0.0027241 | 0.0005767 | -0.27241 | 0.005767 | H | neg | neg
393 | PAMF Pre-COVID | Pre-Pandemic | Negative | 565 |  |  | 136 | 101.3333 | 22404.6667 | 0.0033029 | 0.0012051 | 0.33029 | 0.012051 | H | neg | neg
394 | PAMF Pre-COVID | Pre-Pandemic | Negative | 566 |  |  | 104.6667 | 91 | 19201.5 | 0.002222 | 0.00086799 | 0.2222 | 0.0086799 | H | neg | neg
395 | PAMF Pre-COVID | Pre-Pandemic | Negative | 567 |  |  | 63.6667 | 94 | 21038.6667 | 7.9219e-05 | 0.00093479 | 0.0079219 | 0.0093479 | H | neg | neg
396 | PAMF Pre-COVID | Pre-Pandemic | Negative | 568 |  |  | 185.3333 | 10 | 25660.1667 | 0.0045076 | -0.0024811 | 0.45076 | -0.024811 | H | neg | neg
397 | PAMF Pre-COVID | Pre-Pandemic | Negative | 569 |  |  | -12.1667 | 114 | 27499.5 | -0.0029758 | 0.0014667 | -0.29758 | 0.014667 | H | neg | neg
398 | PAMF Pre-COVID | Pre-Pandemic | Negative | 570 |  |  | -3 | 64.3333 | 24153.1667 | -0.0030086 | -0.00038642 | -0.30086 | -0.0038642 | H | neg | neg
399 | PAMF Pre-COVID | Pre-Pandemic | Negative | 571 |  |  | 106.6667 | 86.6667 | 26860.8333 | 0.0013775 | 0.00048398 | 0.13775 | 0.0048398 | H | neg | neg
400 | PAMF Pre-COVID | Pre-Pandemic | Negative | 572 |  |  | 83.5 | 162.3333 | 23496.3333 | 0.00058874 | 0.0037736 | 0.058874 | 0.037736 | H | neg | neg
401 | PAMF Pre-COVID | Pre-Pandemic | Negative | 573 |  |  | -1.3333 | 108.3333 | 25122.1667 | -0.0028262 | 0.0013799 | -0.28262 | 0.013799 | H | neg | neg
402 | PAMF Pre-COVID | Pre-Pandemic | Negative | 574 |  |  | 41.6667 | 94.1667 | 24337.5 | -0.0011505 | 0.00084232 | -0.11505 | 0.0084232 | H | neg | neg
403 | PAMF Pre-COVID | Pre-Pandemic | Negative | 575 |  |  | 48.5 | 152.5 | 25661.5 | -0.00082484 | 0.003072 | -0.082484 | 0.03072 | H | neg | neg
404 | PAMF Pre-COVID | Pre-Pandemic | Negative | 576 |  |  | 101.1667 | 89 | 23302.5 | 0.0013518 | 0.00065801 | 0.13518 | 0.0065801 | H | neg | neg
405 | PAMF Pre-COVID | Pre-Pandemic | Negative | 577 |  |  | 4.3333 | 111 | 25319.8333 | -0.0025803 | 0.0014745 | -0.25803 | 0.014745 | H | neg | neg
406 | PAMF Pre-COVID | Pre-Pandemic | Negative | 578 |  |  | 106.1667 | 4.6667 | 21104 | 0.0017295 | -0.0032695 | 0.17295 | -0.032695 | H | neg | neg
407 | PAMF Pre-COVID | Pre-Pandemic | Negative | 579 |  |  | 102.5 | 102.3333 | 21519.1667 | 0.0015258 | 0.0013321 | 0.15258 | 0.013321 | H | neg | neg
408 | PAMF Pre-COVID | Pre-Pandemic | Negative | 580 |  |  | -78 | 97.8333 | 31585.8333 | -0.0046593 | 0.00095507 | -0.46593 | 0.0095507 | H | neg | neg
409 | PAMF Pre-COVID | Pre-Pandemic | Negative | 581 |  |  | 95.6667 | 89 | 27141.3333 | 0.00097637 | 0.00078601 | 0.097637 | 0.0078601 | H | neg | neg
410 | PAMF Pre-COVID | Pre-Pandemic | Negative | 582 |  |  | 101 | 163.6667 | 24791.8333 | 0.001284 | 0.0038722 | 0.1284 | 0.038722 | H | neg | neg
411 | PAMF Pre-COVID | Pre-Pandemic | Negative | 583 |  |  | 1 | 67.3333 | 27704.3333 | -0.0024605 | -1.2032e-05 | -0.24605 | -0.00012032 | H | neg | neg
412 | PAMF Pre-COVID | Pre-Pandemic | Negative | 584 |  |  | -0.66667 | 101.3333 | 23645 | -0.0029534 | 0.0014238 | -0.29534 | 0.014238 | H | neg | neg
413 | PAMF Pre-COVID | Pre-Pandemic | Negative | 585 |  |  | 36.3333 | 188.8333 | 25360.5 | -0.0012947 | 0.0047778 | -0.12947 | 0.047778 | H | neg | neg
414 | PAMF Pre-COVID | Pre-Pandemic | Negative | 586 |  |  | -3.6667 | 95.3333 | 25884 | -0.0028138 | 0.0010689 | -0.28138 | 0.010689 | H | neg | neg
415 | PAMF Pre-COVID | Pre-Pandemic | Negative | 587 |  |  | 131 | 91.3333 | 24761.1667 | 0.0024972 | 0.0009558 | 0.24972 | 0.009558 | H | neg | neg
416 | PAMF Pre-COVID | Pre-Pandemic | Negative | 588 |  |  | 104 | 105.6667 | 24160.5 | 0.0014417 | 0.0015728 | 0.14417 | 0.015728 | H | neg | neg
417 | PAMF Pre-COVID | Pre-Pandemic | Negative | 589 |  |  | 97.8333 | 86 | 22928.8333 | 0.0012502 | 0.00079958 | 0.12502 | 0.0079958 | H | neg | neg
418 | PAMF Pre-COVID | Pre-Pandemic | Negative | 590 |  |  | 5.3333 | 101.5 | 21419.5 | -0.0029802 | 0.0015796 | -0.29802 | 0.015796 | H | neg | neg
419 | FL PCR+ | PCR+ | Positive | 201 |  | NA | 44.6667 | 92 | 14492.1667 | 0.0031051 | 0.0014031 | 0.31051 | 0.014031 | P | neg | neg
420 | FL PCR+ | PCR+ | Positive | 202 |  | NA | 4.1667 | 91.6667 | 14946.1667 | 0.00030108 | 0.0013381 | 0.030108 | 0.013381 | P | neg | neg
421 | FL PCR+ | PCR+ | Positive | 203 |  | NA | 676.8333 | 860.5 | 14990.3333 | 0.045174 | 0.052623 | 4.5174 | 0.52623 | P | pos | neg
422 | FL PCR+ | PCR+ | Positive | 204 |  | NA | 606 | 196.3333 | 13142.5 | 0.041798 | 0.0091053 | 4.1798 | 0.091053 | P | pos | neg
423 | FL PCR+ | PCR+ | Positive | 205 |  | NA | 3788 | 37818.8333 | 15959.8333 | 0.2338 | 2.3648 | 23.3795 | 23.6482 | P | pos | pos
424 | FL PCR+ | PCR+ | Positive | 206 |  | NA | 255 | 134.1667 | 14810.1667 | 0.013392 | 0.0038825 | 1.3392 | 0.038825 | P | neg | neg
425 | FL PCR+ | PCR+ | Positive | 207 |  | NA | 1963 | 3572 | 14327.5 | 0.1367 | 0.24389 | 13.6695 | 2.4389 | P | pos | pos
426 | FL PCR+ | PCR+ | Positive | 208 |  | NA | 26.3333 | 121.3333 | 14836.1667 | 0.0014716 | 0.0029433 | 0.14716 | 0.029433 | P | neg | neg
427 | FL PCR+ | PCR+ | Positive | 209 |  | NA | 275.3333 | 1017.3333 | 15087.3333 | 0.017951 | 0.062282 | 1.7951 | 0.62282 | P | pos | neg
428 | FL PCR+ | PCR+ | Positive | 210 |  | NA | 242.1667 | 70.1667 | 15843 | 0.01518 | -0.00019988 | 1.518 | -0.0019988 | P | pos | neg
429 | FL PCR+ | PCR+ | Positive | 211 |  | NA | 2.6667 | 99.3333 | 14317 | 6.9847e-05 | 0.001816 | 0.0069847 | 0.01816 | P | neg | neg
430 | FL PCR+ | PCR+ | Positive | 212 |  | NA | -96.6667 | 105.6667 | 15872.3333 | -0.0061953 | 0.0020371 | -0.61953 | 0.020371 | P | neg | neg
431 | FL PCR+ | PCR+ | Positive | 213 |  | NA | 700 | 371.1667 | 13586 | 0.051401 | 0.021922 | 5.1401 | 0.21922 | P | pos | neg
432 | FL PCR+ | PCR+ | Positive | 214 |  | NA | 0 | 0 | 18066.3333 | -9.2253e-05 | -0.0040591 | -0.0092253 | -0.040591 | P | neg | neg
433 | FL PCR+ | PCR+ | Positive | 215 |  | NA | 103.8333 | 186.6667 | 14863.5 | 0.0068737 | 0.0076249 | 0.68737 | 0.076249 | P | neg | neg
434 | FL PCR+ | PCR+ | Positive | 216 |  | NA | 87 | 99.3333 | 17120.1667 | 0.0049844 | 0.0015187 | 0.49844 | 0.015187 | P | neg | neg
435 | FL PCR+ | PCR+ | Positive | 217 |  | NA | 1240.3333 | 35653.3333 | 13015 | 0.095172 | 2.7338 | 9.5172 | 27.3377 | P | pos | pos
436 | FL PCR+ | PCR+ | Positive | 218 |  | NA | 344.1667 | 109.6667 | 15121 | 0.022651 | 0.0024028 | 2.2651 | 0.024028 | P | pos | neg
437 | FL PCR+ | PCR+ | Positive | 219 |  | NA | 7.3333 | 98.1667 | 14946.5 | 0.00037913 | 0.0016615 | 0.037913 | 0.016615 | P | neg | neg
438 | FL PCR+ | PCR+ | Positive | 220 |  | NA | 272.3333 | 149 | 15024 | 0.018016 | 0.0050364 | 1.8016 | 0.050364 | P | pos | neg
439 | FL PCR+ | PCR+ | Positive | 221 |  | NA | 1870 | 14103.1667 | 14826.3333 | 0.12601 | 0.94628 | 12.6015 | 9.4628 | P | pos | pos
440 | FL PCR+ | PCR+ | Positive | 222 |  | NA | -37.6667 | 627.6667 | 15310 | -0.0025691 | 0.036207 | -0.25691 | 0.36207 | P | neg | neg
441 | FL PCR+ | PCR+ | Positive | 223 |  | NA | -16.3333 | 399 | 15271.6667 | -0.0011787 | 0.021325 | -0.11787 | 0.21325 | P | neg | neg
442 | FL PCR+ | PCR+ | Positive | 441 |  | NA | 82 | 95 | 14994 | -0.0013783 | 0.0007114 | -0.13783 | 0.007114 | P | neg | neg
443 | FL PCR+ | PCR+ | Positive | 442 |  | NA | 21274.6667 | 65294.6667 | 16989.3333 | 1.2462 | 3.8383 | 124.6194 | 38.3831 | P | pos | pos
444 | FL PCR+ | PCR+ | Positive | 443 |  | NA | 8539.5 | 754.3333 | 14235.3333 | 0.59267 | 0.047066 | 59.2668 | 0.47066 | P | pos | neg
445 | FL PCR+ | PCR+ | Positive | 444 |  | NA | 27047 | 65240.6667 | 16036 | 1.6802 | 4.0631 | 168.024 | 40.6313 | P | pos | pos
446 | FL PCR+ | PCR+ | Positive | 445 |  | NA | 91.3333 | 36.5 | 16647.5 | -0.00068078 | -0.0028733 | -0.068078 | -0.028733 | P | neg | neg
447 | FL PCR+ | PCR+ | Positive | 446 |  | NA | 98.1667 | 107.3333 | 14262.1667 | -0.00031552 | 0.0016127 | -0.031552 | 0.016127 | P | neg | neg
448 | FL PCR+ | PCR+ | Positive | 447 |  | NA | 628.3333 | 8017.1667 | 15186.5 | 0.034757 | 0.52295 | 3.4757 | 5.2295 | P | pos | pos
449 | FL PCR+ | PCR+ | Positive | 448 |  | NA | 3418.6667 | 19725.6667 | 13545.8333 | 0.24496 | 1.4507 | 24.4958 | 14.5066 | P | pos | pos
450 | FL PCR+ | PCR+ | Positive | 449 |  | NA | 109.3333 | 165.6667 | 15229.1667 | 0.00058003 | 0.0059316 | 0.058003 | 0.059316 | P | neg | neg
451 | FL PCR+ | PCR+ | Positive | 450 |  | NA | 206.6667 | 101 | 13979.3333 | 0.0075945 | 0.001836 | 0.75945 | 0.01836 | P | neg | neg
452 | FL PCR+ | PCR+ | Positive | 451 |  | NA | 492.3333 | 6834.1667 | 14333.6667 | 0.027337 | 0.47154 | 2.7337 | 4.7154 | P | pos | pos
453 | FL PCR+ | PCR+ | Positive | 452 |  | NA | 10591.6667 | 65240.6667 | 13957.8333 | 0.75163 | 4.6687 | 75.1633 | 46.6873 | P | pos | pos
454 | LL PCR+ | PCR+ | Positive | 13028JL |  | NA | 9355.1667 | 65237 | 14838.8333 | 0.63047 | 4.3915 | 63.0474 | 43.9154 | P | pos | pos
455 | LL PCR+ | PCR+ | Positive | 14902DO |  | NA | 951.3333 | 19335 | 12521.1667 | 0.076005 | 1.5385 | 7.6005 | 15.3846 | P | pos | pos
456 | LL PCR+ | PCR+ | Positive | 18254AV |  | NA | 2665.1667 | 65240.6667 | 15528.8333 | 0.17165 | 4.1966 | 17.1648 | 41.9664 | P | pos | pos
457 | LL PCR+ | PCR+ | Positive | 20455WF |  | NA | 3927 | 65270.3333 | 14740.5 | 0.26643 | 4.4231 | 26.6431 | 44.231 | P | pos | pos
458 | LL PCR+ | PCR+ | Positive | 24897JN |  | NA | 5445.6667 | 65242 | 14961.6667 | 0.364 | 4.3558 | 36.3997 | 43.5582 | P | pos | pos
459 | LL PCR+ | PCR+ | Positive | 256531IM |  | NA | 2763.5 | 9806.3333 | 13083.5 | 0.21125 | 0.74404 | 21.1246 | 7.4404 | P | pos | pos
460 | LL PCR+ | PCR+ | Positive | 26437RM |  | NA | 2812.6667 | 20346.8333 | 14417.8333 | 0.19511 | 1.4063 | 19.5106 | 14.0626 | P | pos | pos
461 | LL PCR+ | PCR+ | Positive | 65363RT |  | NA | 1783.6667 | 5974.6667 | 13111.5 | 0.13606 | 0.45022 | 13.6064 | 4.5022 | P | pos | pos
462 | LL PCR+ | PCR+ | Positive | 91615MB |  | NA | 1590.3333 | 46963 | 15350.8333 | 0.10362 | 3.0546 | 10.3621 | 30.5464 | P | pos | pos
463 | LL PCR+ | PCR+ | Positive | 30622DK |  | NA | 6765.1667 | 65244 | 15302.5 | 0.43839 | 4.2586 | 43.8392 | 42.5861 | P | pos | pos
464 | LL PCR+ | PCR+ | Positive | 32395LP |  | NA | 3448.6667 | 54595.3333 | 14849 | 0.22843 | 3.6715 | 22.8433 | 36.7154 | P | pos | pos
465 | LL PCR+ | PCR+ | Positive | 32540EJ |  | NA | 4413 | 65240.3333 | 16595.6667 | 0.2625 | 3.9265 | 26.2498 | 39.2655 | P | pos | pos
466 | LL PCR+ | PCR+ | Positive | 35264HC |  | NA | 5819 | 61882.3333 | 14725.3333 | 0.39132 | 4.1972 | 39.1321 | 41.9723 | P | pos | pos
467 | LL PCR+ | PCR+ | Positive | 40590DC |  | NA | 43034 | 65322 | 17451 | 2.4627 | 3.7388 | 246.2743 | 37.3877 | P | pos | pos
468 | LL PCR+ | PCR+ | Positive | 44408AO |  | NA | 2919.3333 | 2153.6667 | 13840.5 | 0.20683 | 0.15007 | 20.6833 | 1.5007 | P | pos | neg
469 | LL PCR+ | PCR+ | Positive | 57728ER |  | NA | 967.6667 | 292 | 15563 | 0.058536 | 0.013836 | 5.8536 | 0.13836 | P | pos | neg
470 | LL PCR+ | PCR+ | Positive | 65698LM |  | NA | 8568.8333 | 65245.6667 | 12826.6667 | 0.66363 | 5.0807 | 66.363 | 50.8074 | P | pos | pos
471 | LL PCR+ | PCR+ | Positive | 34005MH |  | NA | 5995.8333 | 5504.1667 | 15086.3333 | 0.39368 | 0.35976 | 39.3679 | 3.5976 | P | pos | pos
472 | LL PCR+ | PCR+ | Positive | 45612JS |  | NA | 11418.3333 | 65302.6667 | 15878.8333 | 0.71881 | 4.1077 | 71.8808 | 41.0767 | P | pos | pos
473 | LL PCR+ | PCR+ | Positive | 90003MS |  | NA | 459.8333 | 469.8333 | 17034.5 | 0.02673 | 0.023022 | 2.673 | 0.23022 | P | pos | neg
474 | LL PCR+ | PCR+ | Positive | 54225BW |  | NA | 4022.8333 | 12200 | 15983.1667 | 0.25141 | 0.75844 | 25.141 | 7.5844 | P | pos | pos
475 | LL PCR+ | PCR+ | Positive | 21531NH |  | NA | 2496.5 | 10171.5 | 17209.8333 | 0.1448 | 0.58652 | 14.4801 | 5.8652 | P | pos | pos
476 | LL PCR+ | PCR+ | Positive | 41279FR |  | NA | 21261.8333 | 65030 | 18239.5 | 1.1655 | 3.5611 | 116.5456 | 35.6108 | P | pos | pos
477 | LL PCR+ | PCR+ | Positive | 89439MC |  | NA | 1725.8333 | 2848.3333 | 15939.8333 | 0.10799 | 0.17382 | 10.7989 | 1.7382 | P | pos | pos
478 | LL PCR+ | PCR+ | Positive | 68775KL |  | NA | 1979.1667 | 998.1667 | 14721.5 | 0.13413 | 0.062528 | 13.4135 | 0.62528 | P | pos | neg
479 | LL PCR+ | PCR+ | Positive | 83211PH |  | NA | 4309.8333 | 65144.6667 | 14214.1667 | 0.30289 | 4.5776 | 30.289 | 45.7762 | P | pos | pos
480 | LL PCR+ | PCR+ | Positive | 93272RW |  | NA | 101 | 73.3333 | 14457.8333 | 0.0066746 | -0.00029972 | 0.66746 | -0.0029972 | P | neg | neg
481 | Cross-Reactivity | Cross-Reactivity | Negative | MPV-7 | AcHBs |  | -0.16667 | 9.6667 | 12027.3333 | -0.0029654 | -0.0052931 | -0.29654 | -0.052931 | H | neg | neg
482 | Cross-Reactivity | Cross-Reactivity | Negative | MPV-9 | AgHBs |  | -5.3333 | 195.3333 | 13169.6667 | -0.0031005 | 0.0092637 | -0.31005 | 0.092637 | H | neg | neg
483 | Cross-Reactivity | Cross-Reactivity | Negative | MPV-10 | AgHBs |  | 2.5 | 107.6667 | 10665 | -0.0030941 | 0.0032195 | -0.30941 | 0.032195 | H | neg | neg
484 | Cross-Reactivity | Cross-Reactivity | Negative | MPV-15 | AgHBs |  | -7.3333 | 102.6667 | 11953 | -0.0035834 | 0.0024543 | -0.35834 | 0.024543 | H | neg | neg
485 | Cross-Reactivity | Cross-Reactivity | Negative | MPV-31 | AgHBs |  | 63 | 67.3333 | 11831.5 | 0.0023242 | -0.00050682 | 0.23242 | -0.0050682 | H | neg | neg
486 | Cross-Reactivity | Cross-Reactivity | Negative | Boca-5725 | Flu A / Flu B |  | -105 | 37.3333 | 12112.5 | -0.0088049 | -0.0032706 | -0.88049 | -0.032706 | H | neg | neg
487 | Cross-Reactivity | Cross-Reactivity | Negative | Boca-5671 | Flu A / Flu B |  | -3.3333 | 43.3333 | 10339.6667 | -0.0004526 | -0.0032511 | -0.04526 | -0.032511 | H | neg | neg
488 | Cross-Reactivity | Cross-Reactivity | Negative | Boca-6304 | Flu A / Flu B |  | 1.3333 | 105.6667 | 12256 | 2.7278e-07 | 0.0023599 | 2.7278e-05 | 0.023599 | H | neg | neg
489 | Cross-Reactivity | Cross-Reactivity | Negative | Boca-5778 | Flu A / Flu B |  | 2.3333 | 74.3333 | 10639.6667 | 9.4624e-05 | -0.00023546 | 0.0094624 | -0.0023546 | H | neg | neg
490 | Cross-Reactivity | Cross-Reactivity | Negative | Boca-7010 | Flu A / Flu B |  | 4 | 104.6667 | 11072 | 0.00024194 | 0.0025224 | 0.024194 | 0.025224 | H | neg | neg
491 | Cross-Reactivity | Cross-Reactivity | Negative | Disc-4255 | Flu A |  | -1 | 20 | 10643.5 | -0.00021966 | -0.0053577 | -0.021966 | -0.053577 | H | neg | neg
492 | Cross-Reactivity | Cross-Reactivity | Negative | Disc-4251 | Flu A |  | 63.6667 | 11.6667 | 10876.1667 | 0.0057507 | -0.0060115 | 0.57507 | -0.060115 | H | neg | neg
493 | Cross-Reactivity | Cross-Reactivity | Negative | Disc-4253 | Flu A |  | 5.3333 | 23 | 10654.5 | 0.00037703 | -0.0050696 | 0.037703 | -0.050696 | H | neg | neg
494 | Cross-Reactivity | Cross-Reactivity | Negative | CDC-0032 | RSV |  | -48.3333 | 98 | 13555.8333 | -0.0061841 | 0.0018198 | -0.61841 | 0.018198 | H | neg | neg
495 | Cross-Reactivity | Cross-Reactivity | Negative | CDC-1524 | RSV |  | -34.3333 | 106 | 9736.1667 | -0.0071723 | 0.0033554 | -0.71723 | 0.033554 | H | neg | neg
496 | Cross-Reactivity | Cross-Reactivity | Negative | CDC-0201 | RSV |  | -61.8333 | 74.1667 | 11163.6667 | -0.0087185 | 7.4943e-05 | -0.87185 | 0.00074943 | H | neg | neg
497 | Cross-Reactivity | Cross-Reactivity | Negative | CDC-0093 | RSV |  | -3.3333 | 42.6667 | 10912 | -0.0035587 | -0.00281 | -0.35587 | -0.0281 | H | neg | neg
498 | Cross-Reactivity | Cross-Reactivity | Negative | CDC-0139 | RSV |  | 0.33333 | 15.1667 | 12650 | -0.0027799 | -0.0045978 | -0.27799 | -0.045978 | H | neg | neg
499 | Cross-Reactivity | Cross-Reactivity | Negative | 1 | ANA |  | -99.3333 | 35.6667 | 12957.5 | -0.012268 | -0.0030601 | -1.2268 | -0.030601 | H | neg | neg
500 | Cross-Reactivity | Cross-Reactivity | Negative | 2 | ANA |  | -59.1667 | 8 | 13253 | -0.0089642 | -0.0050789 | -0.89642 | -0.050789 | H | neg | neg
501 | Cross-Reactivity | Cross-Reactivity | Negative | 3 | ANA |  | -63 | 2167.5 | 13108.3333 | -0.0093554 | 0.15956 | -0.93554 | 1.5956 | H | neg | neg
502 | Cross-Reactivity | Cross-Reactivity | Negative | 4 | ANA |  | -100.3333 | 100 | 12925.6667 | -0.012375 | 0.001908 | -1.2375 | 0.01908 | H | neg | neg
503 | Cross-Reactivity | Cross-Reactivity | Negative | 5 | ANA |  | 2.8333 | 61.8333 | 13078 | -0.0043447 | -0.0010317 | -0.43447 | -0.010317 | H | neg | neg
504 | Cross-Reactivity | Cross-Reactivity | Negative | 6 | ANA |  | -5.6667 | 99.8333 | 11598.5 | -0.0056313 | 0.0021119 | -0.56313 | 0.021119 | H | neg | neg
505 | Cross-Reactivity | Cross-Reactivity | Negative | 7 | ANA |  | -61.3333 | 101.8333 | 12700.3333 | -0.0095247 | 0.0020862 | -0.95247 | 0.020862 | H | neg | neg
506 | Cross-Reactivity | Cross-Reactivity | Negative | 8 | ANA |  | 3.6667 | 12.6667 | 11156 | -0.0050183 | -0.0056151 | -0.50183 | -0.056151 | H | neg | neg
507 | Cross-Reactivity | Cross-Reactivity | Negative | 9 | ANA |  | -67.6667 | 100.3333 | 11702.5 | -0.010878 | 0.0021359 | -1.0878 | 0.021359 | H | neg | neg
508 | Cross-Reactivity | Cross-Reactivity | Negative | 10 | ANA |  | 2.3333 | 100.6667 | 11035 | -0.0051941 | 0.0022952 | -0.51941 | 0.022952 | H | neg | neg
509 | Cross-Reactivity | Cross-Reactivity | Negative | 11 | ANA |  | -69 | 42.1667 | 10912.3333 | -0.0064665 | -0.0031871 | -0.64665 | -0.031871 | H | neg | neg
510 | Cross-Reactivity | Cross-Reactivity | Negative | 12 | ANA |  | -38 | 39 | 10487.6667 | -0.0037632 | -0.0036196 | -0.37632 | -0.036196 | H | neg | neg
511 | Cross-Reactivity | Cross-Reactivity | Negative | 0123 V-1 | ZIKV |  | 37.5 | 106 | 17072 | 0.0019135 | 0.0020599 | 0.19135 | 0.020599 | H | neg | neg
512 | Cross-Reactivity | Cross-Reactivity | Negative | 0123 V-2 | ZIKV |  | -3 | 103 | 17544.1667 | 0.001691 | -0.00030399 | 0.1691 | -0.0030399 | H | neg | neg
513 | Cross-Reactivity | Cross-Reactivity | Negative | 0123 V-4 | ZIKV |  | -1.8333 | 103 | 14850 | 0.0019978 | -0.00028058 | 0.19978 | -0.0028058 | H | neg | neg
514 | Cross-Reactivity | Cross-Reactivity | Negative | 0123 V-5 | ZIKV |  | 4.3333 | 107.6667 | 15259.5 | 0.00225 | 0.00013107 | 0.225 | 0.0013107 | H | neg | neg
515 | Cross-Reactivity | Cross-Reactivity | Negative | 0123 V-6 | ZIKV |  | -108.6667 | 97.5 | 16893.3333 | 0.0013911 | -0.006472 | 0.13911 | -0.06472 | H | neg | neg
516 | Cross-Reactivity | Cross-Reactivity | Negative | 0123 V-7 | ZIKV |  | -7.6667 | 12.3333 | 15036 | -0.0041013 | -0.00055423 | -0.41013 | -0.0055423 | H | neg | neg
517 | Cross-Reactivity | Cross-Reactivity | Negative | 0123 V-8 | ZIKV |  | -69 | 97.6667 | 14779.5 | 0.0016013 | -0.0047137 | 0.16013 | -0.047137 | H | neg | neg
518 | Cross-Reactivity | Cross-Reactivity | Negative | 150-1 | ZIKV |  | -138.8333 | 100 | 19792.6667 | 0.0013473 | -0.0071323 | 0.13473 | -0.071323 | H | neg | neg
519 | Cross-Reactivity | Cross-Reactivity | Negative | 150-2 | ZIKV |  | -32.8333 | 7 | 16577.5 | -0.0040014 | -0.0021213 | -0.40014 | -0.021213 | H | neg | neg
520 | Cross-Reactivity | Cross-Reactivity | Negative | 150-3 | ZIKV |  | -65.6667 | 2 | 17244.5 | -0.0041366 | -0.0039433 | -0.41366 | -0.039433 | H | neg | neg
521 | Cross-Reactivity | Cross-Reactivity | Negative | 150-4 | ZIKV |  | -109.3333 | 9.6667 | 14560 | -0.0043727 | -0.0076694 | -0.43727 | -0.076694 | H | neg | neg
522 | Cross-Reactivity | Cross-Reactivity | Negative | 150-5 | ZIKV |  | 7 | 77 | 17610.6667 | 0.00020821 | 0.00026499 | 0.020821 | 0.0026499 | H | neg | neg
523 | Cross-Reactivity | Cross-Reactivity | Negative | 150-6 | ZIKV |  | 2 | 100 | 15880.1667 | 0.0016792 | -2.0991e-05 | 0.16792 | -0.00020991 | H | neg | neg
524 | Cross-Reactivity | Cross-Reactivity | Negative | 150-7 | ZIKV |  | 2 | 106.1667 | 17607 | 0.0018648 | -1.8932e-05 | 0.18648 | -0.00018932 | H | neg | neg
525 | Cross-Reactivity | Cross-Reactivity | Negative | 150-8 | ZIKV |  | 6.6667 | 76.6667 | 17322 | 0.00019243 | 0.00025016 | 0.019243 | 0.0025016 | H | neg | neg
526 | Cross-Reactivity | Cross-Reactivity | Negative | 246 | HIV |  | 4.3333 | 96.3333 | 12390.3333 | 0.00033628 | 0.0037798 | 0.033628 | 0.037798 | H | neg | neg
527 | Cross-Reactivity | Cross-Reactivity | Negative | coV229E-1 | Coronavirus 229E |  | -6.6667 | 75.6667 | 15910 | -0.0042007 | 4.1902e-05 | -0.42007 | 0.00041902 | H | neg | neg
528 | Cross-Reactivity | Cross-Reactivity | Negative | HCV-1 | HCV |  | 0 | 167 | 13998.6667 | -0.004298 | 0.0065721 | -0.4298 | 0.065721 | H | neg | neg
529 | Cross-Reactivity | Cross-Reactivity | Negative | 05571-CA | ANA |  | -35 | 32.6667 | 10014 | 0.0039611 | -0.0037947 | 0.39611 | -0.037947 | H | neg | neg
530 | Cross-Reactivity | Cross-Reactivity | Negative | 63311-LC | Coronavirus OC43 |  | -92 | 69.5 | 10880.6667 | -0.001593 | -0.00010722 | -0.1593 | -0.0010722 | H | neg | neg
531 | Cross-Reactivity | Cross-Reactivity | Negative | 77294-JM | Coronavirus NL63 |  | -1.3333 | 40 | 10317.3333 | 0.0071078 | -0.0029723 | 0.71078 | -0.029723 | H | neg | neg
532 | Cross-Reactivity | Cross-Reactivity | Negative | 91466-JP | Haemophilus influenzae |  | -1 | 9.6667 | 9975.5 | 0.0073848 | -0.006115 | 0.73848 | -0.06115 | H | neg | neg
533 | Cross-Reactivity | Cross-Reactivity | Negative | 89230-EC | Rheumatoid Factor |  | -10.1667 | 79.3333 | 9787.8333 | 0.0065898 | 0.00088545 | 0.65898 | 0.0088545 | H | neg | neg
534 | Cross-Reactivity | Cross-Reactivity | Negative | 59565-DD | RSV |  | 0.66667 | 34.8333 | 10607.1667 | 0.0071021 | -0.0033782 | 0.71021 | -0.033782 | H | neg | neg
535 | Cross-Reactivity | Cross-Reactivity | Negative | 05646-RK | RSV |  | -1.3333 | 89.1667 | 9708.3333 | 0.0075536 | 0.0019056 | 0.75536 | 0.019056 | H | neg | neg
536 | Cross-Reactivity | Cross-Reactivity | Negative | 74985-AL | RSV |  | 9.3333 | 102.8333 | 9584.3333 | 0.0087643 | 0.0033562 | 0.87643 | 0.033562 | H | neg | neg
537 | Cross-Reactivity | Cross-Reactivity | Negative | 591 | HIV |  | 5.6667 | 64.3333 | 16613.1667 | 0.003732 | -0.0044543 | 0.3732 | -0.044543 | H | neg | neg
538 | Cross-Reactivity | Cross-Reactivity | Negative | 592 | HIV |  | 104.6667 | 67.1667 | 13310.1667 | 0.004871 | 0.0018783 | 0.4871 | 0.018783 | H | neg | neg
539 | Cross-Reactivity | Cross-Reactivity | Negative | 593 | HIV |  | 95 | 2.3333 | 16857.5 | 0 | 0.00090959 | 0 | 0.0090959 | H | neg | neg
540 | Cross-Reactivity | Cross-Reactivity | Negative | 594 | HIV |  | 95 | 31.8333 | 16084.8333 | 0.001834 | 0.00095328 | 0.1834 | 0.0095328 | H | neg | neg
541 | Cross-Reactivity | Cross-Reactivity | Negative | 595 | HIV |  | 165.5 | -5 | 18196 | -0.00040302 | 0.0047172 | -0.040302 | 0.047172 | H | neg | neg
542 | Cross-Reactivity | Cross-Reactivity | Negative | 596 | HIV |  | 48.3333 | 35.1667 | 16210.8333 | 0.0020254 | -0.0019329 | 0.20254 | -0.019329 | H | neg | neg
543 | Cross-Reactivity | Cross-Reactivity | Negative | 597 | HIV |  | 89 | -30.3333 | 15189 | -0.0021507 | 0.00061448 | -0.21507 | 0.0061448 | H | neg | neg
544 | Cross-Reactivity | Cross-Reactivity | Negative | 598 | HIV |  | 135.1667 | -7.5 | 13949.1667 | -0.00070494 | 0.0039787 | -0.070494 | 0.039787 | H | neg | neg
545 | Cross-Reactivity | Cross-Reactivity | Negative | 599 | HIV |  | 87.6667 | -0.5 | 16190.1667 | -0.000175 | 0.00049413 | -0.0175 | 0.0049413 | H | neg | neg
546 | Cross-Reactivity | Cross-Reactivity | Negative | 7166 | HCV |  | 39.3333 | 215.3333 | 18671.1667 | 0.0018924 | 0.0074982 | 0.18924 | 0.074982 | H | neg | neg
547 | Cross-Reactivity | Cross-Reactivity | Negative | 7318 | HCV |  | 181 | 62.5 | 16469.1667 | 0.010747 | -0.00077923 | 1.0747 | -0.0077923 | H | neg | neg
548 | Cross-Reactivity | Cross-Reactivity | Negative | 7180 | HCV |  | 47 | 102.5 | 17351.6667 | 0.0024781 | 0.0015657 | 0.24781 | 0.015657 | H | neg | neg
549 | Cross-Reactivity | Cross-Reactivity | Negative | 7612 | HCV |  | -0.16667 | 29.3333 | 13877.5 | -0.0050321 | -0.0034348 | -0.50321 | -0.034348 | H | neg | neg
550 | Cross-Reactivity | Cross-Reactivity | Negative | 7614 | HCV |  | 90.3333 | 107.6667 | 14543.1667 | 0.0014211 | 0.0021087 | 0.14211 | 0.021087 | H | neg | neg
551 | In-House Healthy | Pre-Pandemic | Negative | 53 |  |  | 91.3333 | 102.1667 | 13864.6667 | 0.0019234 | 0.0019113 | 0.19234 | 0.019113 | H | neg | neg
552 | In-House Healthy | Pre-Pandemic | Negative | 54 |  |  | 129.3333 | 17.3333 | 12858.8333 | 0.005029 | -0.0045364 | 0.5029 | -0.045364 | H | neg | neg
553 | In-House Healthy | Pre-Pandemic | Negative | 55 |  |  | 101.6667 | 114 | 13935.3333 | 0.004497 | 0.0025594 | 0.4497 | 0.025594 | H | neg | neg
554 | In-House Healthy | Pre-Pandemic | Negative | 56 |  |  | 95.3333 | 106 | 14598.1667 | 0.0038589 | 0.0018952 | 0.38589 | 0.018952 | H | neg | neg
555 | In-House Healthy | Pre-Pandemic | Negative | 57 |  |  | 154.5 | 76.6667 | 13558 | 0.008519 | -0.00012293 | 0.8519 | -0.0012293 | H | neg | neg
556 | In-House Healthy | Pre-Pandemic | Negative | 58 |  |  | 74.3333 | 69.6667 | 14368 | 0.0024592 | -0.00060319 | 0.24592 | -0.0060319 | H | neg | neg
557 | In-House Healthy | Pre-Pandemic | Negative | 59 |  |  | 97.1667 | 112.6667 | 12990.3333 | 0.0044777 | 0.002643 | 0.44777 | 0.02643 | H | neg | neg
558 | In-House Healthy | Pre-Pandemic | Negative | 60 |  |  | 103.8333 | 110 | 13196.8333 | 0.0044645 | 0.0010987 | 0.44645 | 0.010987 | H | neg | neg
559 | In-House Healthy | Pre-Pandemic | Negative | 61 |  |  | 92 | 75.6667 | 13196.8333 | 0.0044645 | 0.0010987 | 0.44645 | 0.010987 | H | neg | neg
560 | In-House Healthy | Pre-Pandemic | Negative | 62 |  |  | 96.3333 | 73 | 12993.1667 | 0.0044126 | -0.00041047 | 0.44126 | -0.0041047 | H | neg | neg
561 | In-House Healthy | Pre-Pandemic | Negative | 63 |  |  | 93.6667 | 75 | 11573.6667 | 0.0047234 | -0.00028801 | 0.47234 | -0.0028801 | H | neg | neg
562 | In-House Healthy | Pre-Pandemic | Negative | 64 |  |  | 39.6667 | 103.6667 | 12755.5 | 5.2265e-05 | 0.0019861 | 0.0052265 | 0.019861 | H | neg | neg
563 | In-House Healthy | Pre-Pandemic | Negative | 65 |  |  | -227.3333 | 90.8333 | 12019 | -0.022159 | 0.00104 | -2.2159 | 0.0104 | H | neg | neg
564 | In-House Healthy | Pre-Pandemic | Negative | 468 |  |  | 100 | 105.6667 | 16925.3333 | 0.00084686 | 0.0016346 | 0.084686 | 0.016346 | H | neg | neg
565 | In-House Healthy | Pre-Pandemic | Negative | 469 |  |  | 41.1667 | 112.6667 | 16165.6667 | -0.0027527 | 0.0021445 | -0.27527 | 0.021445 | H | neg | neg
566 | In-House Healthy | Pre-Pandemic | Negative | 470 |  |  | 94.3333 | 105.6667 | 16775.8333 | 0.00051662 | 0.0016492 | 0.051662 | 0.016492 | H | neg | neg
567 | In-House Healthy | Pre-Pandemic | Negative | 472 |  |  | 188.3333 | 133.6667 | 17102.1667 | 0.0060031 | 0.0032549 | 0.60031 | 0.032549 | H | neg | neg
568 | In-House Healthy | Pre-Pandemic | Negative | 474 |  |  | 127 | 125.6667 | 18291.1667 | 0.0022597 | 0.002606 | 0.22597 | 0.02606 | H | neg | neg
569 | In-House Healthy | Pre-Pandemic | Negative | 475 |  |  | 95 | 80.6667 | 17888 | 0.00052177 | 0.00014908 | 0.052177 | 0.0014908 | H | neg | neg
570 | In-House Healthy | Pre-Pandemic | Negative | 476 |  |  | 101 | 119 | 17466.1667 | 0.00087789 | 0.0023474 | 0.087789 | 0.023474 | H | neg | neg
571 | In-House Healthy | Pre-Pandemic | Negative | 477 |  |  | 202 | 8.6667 | 15490 | 0.0075102 | -0.004476 | 0.75102 | -0.04476 | H | neg | neg
572 | In-House Healthy | Pre-Pandemic | Negative | 479 |  |  | 99.6667 | 73.6667 | 17196.8333 | 0.0008141 | -0.00025198 | 0.08141 | -0.0025198 | H | neg | neg
573 | In-House Healthy | Pre-Pandemic | Negative | 480 |  |  | 130.6667 | 72.6667 | 16850.8333 | 0.0026705 | -0.0003165 | 0.26705 | -0.003165 | H | neg | neg
574 | In-House Healthy | Pre-Pandemic | Negative | 481 |  |  | 98 | 99 | 15120.5 | 0.00081567 | 0.0013888 | 0.081567 | 0.013888 | H | neg | neg
575 | In-House Healthy | Pre-Pandemic | Negative | 485 |  |  | 154.8333 | 73 | 14973.5 | 0.0046193 | -0.00033392 | 0.46193 | -0.0033392 | H | neg | neg
576 | In-House Healthy | Pre-Pandemic | Negative | 487 |  |  | 68.3333 | 204.1667 | 13101.5 | 0.0022389 | 0.0096045 | 0.22389 | 0.096045 | H | neg | neg
577 | In-House Healthy | Pre-Pandemic | Negative | 488 |  |  | 102.6667 | 109.6667 | 12381.6667 | 0.005142 | 0.0025306 | 0.5142 | 0.025306 | H | neg | neg
578 | In-House Healthy | Pre-Pandemic | Negative | 492 |  |  | -3.3333 | 99.3333 | 12936 | -0.0032725 | 0.0016234 | -0.32725 | 0.016234 | H | neg | neg
579 | In-House Healthy | Pre-Pandemic | Negative | 493 |  |  | 92.6667 | 108.1667 | 16350.6667 | 0.0017125 | 0.0019877 | 0.17125 | 0.019877 | H | neg | neg
580 | In-House Healthy | Pre-Pandemic | Negative | 495 |  |  | 193.3333 | 103.8333 | 15259.5 | 0.0084319 | 0.0018458 | 0.84319 | 0.018458 | H | neg | neg
581 | In-House Healthy | Pre-Pandemic | Negative | 498 |  |  | 8.6667 | 103.3333 | 14628.3333 | -0.0038282 | 0.0018913 | -0.38282 | 0.018913 | H | neg | neg
582 | In-House Healthy | Pre-Pandemic | Negative | 499 |  |  | 102.3333 | 103.6667 | 15909.3333 | 0.0023676 | 0.00176 | 0.23676 | 0.0176 | H | neg | neg
583 | In-House Healthy | Pre-Pandemic | Negative | 500 |  |  | 100.1667 | 197.3333 | 14840.5 | 0.0023921 | 0.0081983 | 0.23921 | 0.081983 | H | neg | neg
584 | In-House Healthy | Pre-Pandemic | Negative | 501 |  |  | 115 | 103.5 | 14346.3333 | 0.0035084 | 0.0019401 | 0.35084 | 0.019401 | H | neg | neg
585 | In-House Healthy | Pre-Pandemic | Negative | 505 |  |  | 66.6667 | 73.5 | 13757.6667 | 0.00014537 | -0.00015749 | 0.014537 | -0.0015749 | H | neg | neg
586 | In-House Healthy | Pre-Pandemic | Negative | 507 |  |  | 180.6667 | 57.6667 | 14727.6667 | 0.0078763 | -0.0012222 | 0.78763 | -0.012222 | H | neg | neg
587 | In-House Healthy | Pre-Pandemic | Negative | 508 |  |  | 96.3333 | 96.3333 | 13926.6667 | 0.0022738 | 0.001484 | 0.22738 | 0.01484 | H | neg | neg
588 | In-House Healthy | Pre-Pandemic | Negative | 510 |  |  | 105 | 106 | 12921.5 | 0.0031214 | 0.0023475 | 0.31214 | 0.023475 | H | neg | neg
589 | In-House Healthy | Pre-Pandemic | Negative | 551 |  |  | 65.6667 | 11 | 12752.5 | 7.8416e-05 | -0.0050709 | 0.0078416 | -0.050709 | H | neg | neg
590 | In-House Healthy | Pre-Pandemic | Negative | 552 |  |  | -85 | 13 | 12302.3333 | -0.012166 | -0.0050939 | -1.2166 | -0.050939 | H | neg | neg
591 | Cross-Reactivity | Cross-Reactivity | Negative | 854 | Mycoplasma pneumoniae |  | 76 | 39.3333 | 12592 |  |  | 0.0060356 | 0.0031237 | H | neg | neg
592 | Cross-Reactivity | Cross-Reactivity | Negative | 1408 | Mycoplasma pneumoniae |  | -0.66667 | 2.8333 | 13520.1667 |  |  | -4.9309e-05 | 0.00020956 | H | neg | neg
593 | Cross-Reactivity | Cross-Reactivity | Negative | 4409 | Mycoplasma pneumoniae |  | -99.8333 | 106.1667 | 13518.6667 |  |  | -0.0073849 | 0.0078533 | H | neg | neg